## hexapod summary
[8 empty] | Camera Hexapod Re-verification 2021 |  |  | Camera Hexapod Re-verification 2022
[8 empty] | (no load on Hexapod&Rotator unit) |  |  | (with ComCam mounted on Rotator&Hexapod)
LSST Camera Hexapod Requirement specifications re-verification. Ordered from tighter to looser specs, based on document LTS-206, [1] | [10 empty] | February-March 2022 w/ComCom on Rotator
Specs | ITEM | MOVE | AXIS | MODE | SPEC | ERROR SPEC | MEASUREMENT | MEASURED | MEET SPEC | Notes & | MEASURED | MEET SPEC | Notes &
Index |  |  |  |  | (LTS-206) |  | Instrument | ERROR |  | Comments | ERROR |  | Comments
1 | Translation Resolution in Z (Focus) (LTS-206, 3.3.7) | Trans | Z | Offset | 1um | <50% of move | Linear Gauge | 21% | YES |  | ~40% | YES (marginal) (**) | does the 1um offset 3 out of 5 moves (offsets of 2um always work!)
2 | Rotational Resolution Rx and Ry (Tip-Tilt) (LTS-206, 3.3.9) | Rot | Rx | Offset | 8.19x10-5Deg | <50% of move | Linear Gauge | 13% | YES |  | ~42% | YES | one outlier error >50%
|  |  | Ry | Offset | 8.19x10-5Deg | <50% of move | Linear Gauge | 25% | YES |  | ~44% | YES | two outlier errors >50%
3 | Translation Resolution in X and Y (Decentre) (LTS-602, 3.3.5) | Trans | X | Offset | 5um | <50% of move | Linear Gauge | 35% | YES |  | 12% | YES | always does the offset
|  |  | Y | Offset | 5um | <50% of move | Linear Gauge | 35% | YES |  | 16% | YES | always does the offset
4 | Rotational Resolution Rz (hexapod twist) (LTS-602, 3.3.11) | Rot | Rz | Offset | 60X10-5Deg | <50% of move | Linear Gauge | 32% | YES |  | Not done | ------ | low priority
5 | Hexapod Repeatability (Operational requirement) (LTS-602, 3.3.12) | Trans | Z | Absolute | +/-50um | <4um | Linear Gauge | <3um | YES |  | <4um | YES
| (steps of 50um in translation, tb/reached to <4xResolution) |  | X | Absolute | +/-50um | <20um | Linear Gauge | <3um | YES |  | <1um | YES
|  |  | Y | Absolute | +/-50um | <20um | Linear Gauge | <4um | YES |  | <2um | YES
| (steps of 0.0015Deg in rotation,tb/reached to <4xResolution) | Rot | Rx | Absolute | +/-0.0015Deg | <0.000328Deg | Linear Gauge | <0.00018Deg | YES |  | <8.1E-5Deg | YES
|  |  | Ry | Absolute | +/-0.0015Deg | <0.000328Deg | Linear Gauge | <0.00018Deg | YES |  | <33.0E-5Deg | YES
|  |  | Rz | Absolute | +/-0.0015Deg | <0.0024Deg | Linear Gauge | <0.00026Deg | YES |  | <44.0E-5Deg | YES
6 | Hexapod Absolute Accuracy (hexapod operation spec) (LTS-206, 3.3.13) | Trans | Z | Absolute | +/-7.6mm | <25um | Laser Tracker | 0um | YES (lucky) | 30um x-talk | < 24um | YES
| (steps of 25%,50%,75%,100% of max range, tb/reached to <25xResolution) |  | X | Absolute | +/-7.6mm | <125um | Laser Tracker | <30um | YES |  | < 40um | YES
|  |  | Y | Absolute | +/-7.6mm | <125um | Laser Tracker | <120um | YES |  | < 51um | YES
|  | Rot | Rx | Absolute | +/-0.24Deg | <0.00205Deg | Laser Tracker | <0.0017Deg | YES |  | < 0.0006Deg | YES | doing frequent zeroing
|  |  | Ry | Absolute | +/-0.24Deg | <0.00205Deg | Laser Tracker | <0.0023Deg | YES (Marginal) |  | <0.0003Deg | YES | has improved the data
|  |  | Rz | Absolute | +/-0.1Deg | <0.015Deg | Laser Tracker | Not done | ------ | low priority | Not done | ------ | low priority
7 | Hexapod Translation Z range (hexapod operation spec) (LTS-206, 3.3.6) | Trans | Z | Absolute | +/-7.6mm * | <25um | Laser Tracker | 0um | YES | from 3.3.13 | < 25um | YES
(*where max. Z range should have been +/-8.7mm)
8 | Hexapod radial XY translational range (hexapod operation spec) (LTS-206, 3.3.4) | Trans | XY | Absolute | +/-7.6mm | <125um | Laser Tracker | <90um | YES |  | <32um | YES |  | the detailed results show that the max error is 32 um.
(where XY=(x2+y2)1/2 )
9 | Hexapod Rotational Rx Ry range (hexapod operation spec) (LTS-206, 3.3.8) |  | RxRy | Absolute | +/-0.24Deg | <0.00205Deg | Laser Tracker | <0.0021Deg | YES |  | <0.00070 | YES
(where RxRy= (Rx2+Ry2)1/2 )
10 | Hexapod Positioning (32 positions) (LTS-206, 3.3.1) | Trans | XY | Absolute | +/-7.6mm | <125um | Laser Tracker | <29um RMS | YES |  | max. <119um | YES | stdev 36um
|  |  | Z | Absolute | +/-7.73mm | <25um | Laser Tracker | <31um RMS | YES (Marginal) | L.T. resol. | max.<30um | YES (Marginal) | L.T. res.~30um (3 outliers)
|  | Rot | RxRy | Absolute | +/-0.17Deg | <0.00205Deg | Laser Tracker | <0.0016Deg RMS | YES |  | max.<0.0025Deg | YES (Marginal) | 6 outliers (due to noise)
|  |  | Rz | Absolute | 0 | <0.015Deg | Laser Tracker | Not checked | ------ | tbd | max.<0.004Deg | YES
11 | Hexapod centers of rotation (LTS-206, 3.3.2) | Rot & | Rx | Absolute | 0.04 to 0.24 Deg
| (Y offset measured-expected, due to Rx rotation for a given COR, should be <125um) | Trans | Y |  | Y=pivot*Rx | ΔY<125um | Laser Tracker | <35um RMS | YES |  | max.<19um | YES
12 | Hexapod cross-talk motion (LTS-206, 3.3.3) | Trans | X | Absolute | 0mm | <125um | Laser Tracker | <30um | YES |  | < 50um | YES
| (moves should be reached with inactive axes residuals X and Y<125um and Z<25um) |  | Y | Absolute | 0mm | <125um | Laser Tracker | <50um | YES |  | < 17um | YES
| (data from the measurements of 3.3.4 and 3.3.8 above) |  | Z | Absolute | 0mm | <25um | Laser Tracker | <27um RMS | YES |  | < 13um | YES
|  |  | Rx | Absolute | 0Deg | <0.00205Deg | Laser Tracker | Not checked | ------ |  | max.<0.0037Deg | Yes (marginal) | Stdev =0.0020Deg
|  |  | Ry | Absolute | 0Deg | <0.00205Deg | Laser Tracker | Not checked | ------ |  | <0.0017Deg | YES
|  |  | Rz | Absolute | 0Deg | <0.015Deg | Laser Tracker | Not checked | ------ |  | <0.0049Deg | YES
13 | Hexapod Rotational Rz range (good to have for engineering) (LTS-206, 3.3.10) | Rot | Rz | Absolute | 0.1Deg | <0.015Deg | Laser Tracker | Not checked | ------ | tbd | Not checked | ------
(**)Note: In both tests 1 and 2, the few measurements
not meeting specs, could be due to the gauges setup
(x2+y2)1/2
(Rx2+Ry2)1/2

## list Moog 3.3.1_+other_07062021
MOOG | Camera Hexapod tests (at El=0Deg) | [11 empty] | (offset due to rotation= |  |  | 5.7502 | mm) |  |  |  | pos.error |  | XY error | XY error |  |  | Error | Error | Other data processing &similar | [49 empty] | positioning error |  |  | XY error | XY error |  |  | MOOG | Camera Hexapod tests (at El=0Deg) |  |  |  |  | (offset due to rotation= |  |  | 5.7502 | mm | XY error | XY error |  | Z | Z error | Error | Error |  | MOOG | Camera Hexapod tests (at El=0Deg) |  |  |  |  | (offset due to rotation= |  |  | 5.7502 | mm | XY error | Z error | Error
Hex | COR(-2758400um from base) (0,0,-1938mm, from rotator surface) unless specified otherwise | [5 empty] | Linear gauges data M2 edge, reduced (mm, Deg) | [5 empty] | laser tracker measured data(at Cam.hex&rot. center)(mm & Deg) | [6 empty] | other |  |  |  | (offset | Z error |  | Lin.Gauges | Laser Tracker | comments |  |  | faults record | [44 empty] | other | Xerror | X offset | Yerror | Y offset |  | (offset | Zerror |  | Hex | COR(-2758400um from base) (0,0,-1938mm, from rotator surface) unless specified otherwise | [5 empty] | laser tracker measured data(at Cam.hex&rot. center)(mm & Deg) |  |  |  | (x2+ y2)1/2 | offset |  | error | offset | Laser Tracker | Laser Tracker |  | Hex | COR(-2758400um from base) (0,0,-1938mm, from rotator surface) unless specified otherwise | [5 empty] | laser tracker measured data(at Cam.hex&rot. center)(mm & Deg) |  |  |  | (offset | (offset | Angular
Rot | Commanded moves to position, from M2 hexapod EUI | [14 empty] | plane-to-plane |  |  |  | Calcs. | Xerror | Yerror |  | removed) |  |  | (meas.-set) | (meas.-set) | e.g. required specs to compare |  |  | and other operational comments | [44 empty] | Calcs. |  | removed |  | removed |  | removed) |  |  | Rot | Commanded moves to position, from M2 hexapod EUI | [6 empty] | mm |  | plane-to-plane |  | removed |  |  | removed | (meas.-set) | offset |  | Rot | Commanded moves to position, from M2 hexapod EUI | [6 empty] | mm |  | plane-to-plane | removed) | removed) | (offset
testing | X | Y | Z | Rx | Ry | Rz | X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Rot.ang.¦ | Rx | Ry | Rz |  | mm | mm |  | mm | mm |  | mm or deg | Angle (Deg) | [49 empty] | mm |  | mm |  |  | mm | mm |  | testing | X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Rot.ang.¦ | mm | mm |  | mm | mm | Angle | removed |  | testing | X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Rot.ang.¦ | mm | mm | removed)
3.3.1. | Commanded moves to position, from M2 hexapod EUI X: 07/06/2021 Positioning Tests | [17 empty] | angle error | [52 empty] | 3.3.1. | 07/06/2021 Positioning Tests | [8 empty] | Deg | [5 empty] | Deg | Deg |  | 3.3.1. | 07/06/2021 Positioning Tests | [8 empty] | Deg |  |  | Deg
zero | 0 | 0 | 0 | 0 | 0 | 0 | [6 empty] | 0 | 0 | 0 | [13 empty] | to keep in mind: |  |  |  | displacement offset | [52 empty] | zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [9 empty] | zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5.66 | 0 | 7.73 | 0.17 | 0 | 0 | 5.66 | -5.01 | 7.74 | 0.173 | (this is a repeat of |  | 5.66 | -5.64 | 7.69 | 0.173 |  |  |  |  | 0 | 0.11018 | 0.11018 | 0.010557 | -0.04 |  |  | 0.003 | with SMR or Gauges on Rotator |  |  | 0.17 | -5.64 | 0.16674 | -0.0032573 | [44 empty] | 0 | 0.022 | 0.11018 | 0.073175 | 0.076411 | -0.017589 | -0.04 |  | 1 | 5.66 | 0 | 7.73 | 0.17 | 0 | 0 | 5.66 | -5.64 | 7.69 | 0.173 | 0.076411 | -0.017589 | 0.00030938 | -0.04 | -0.022 | 0.003 | 0.0007 |  | 1 | 5.66 | 0 | 7.73 | 0.17 | 0 | 0 | 5.66 | -5.64 | 7.69 | 0.173 | -0.017589 | -0.022 | 0.0007
2 | 5.66 | 0 | 7.73 | -0.17 | 0 | 0 | 5.64 | 6.33 | 7.79 | 0.1702 | measurements at right) |  | 5.64 | 5.7 | 7.74 | 0.1702 |  |  |  |  | -0.02 | -0.050175 | 0.054014 | -0.045604 | 0.01 |  |  | 0.0002 | if COR<0: |  |  | -0.17 | 5.7 | Measured Angle error between different coordinate references | [44 empty] | -0.02 | -0.003 | -0.050175 | -0.084175 | 0.084229 | -0.0097714 | 0.01 |  | 2 | 5.66 | 0 | 7.73 | -0.17 | 0 | 0 | 5.64 | 5.7 | 7.74 | 0.1702 | 0.087198 | -0.0068019 | 4.6265e-05 | 0.01 | 0.028 | 0.0002 | -0.0021 |  | 2 | 5.66 | 0 | 7.73 | -0.17 | 0 | 0 | 5.64 | 5.7 | 7.74 | 0.1702 | -0.0068019 | 0.028 | -0.0021
3 | 5.66 | 0 | 7.73 | 0 | 0.17 | 0 | [6 empty] | 11.29 | 0.04 | 7.7 | 0.1663 |  |  |  |  | -0.12018 | 0.04 | 0.12666 | 0.027039 | -0.03 |  |  | 0.0037 | Rx+=> displacement -Y |  |  |  |  | 0.014Deg | [44 empty] | -0.12018 | -0.10318 | 0.04 | 0.006 | 0.10335 | 0.0093495 | -0.03 |  | 3 | 5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 11.29 | 0.04 | 7.7 | 0.1663 | 0.098221 | 0.004221 | 1.7817e-05 | -0.03 | -0.012 | 0.0037 | 0.0014 |  | 3 | 5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 11.29 | 0.04 | 7.7 | 0.1663 | 0.004221 | -0.012 | 0.0014
4 | 5.66 | 0 | 7.73 | 0 | -0.17 | 0 | [6 empty] | -0.04 | 0.03 | 7.72 | 0.1707 |  |  |  |  | 0.050175 | 0.03 | 0.05846 | -0.041158 | -0.01 |  |  | 0.0007 | Ry+=> displacement +X | [48 empty] | 0.050175 | 0.067175 | 0.03 | -0.004 | 0.067294 | -0.026706 | -0.01 |  | 4 | 5.66 | 0 | 7.73 | 0 | -0.17 | 0 | -0.04 | 0.03 | 7.72 | 0.1707 | 0.072514 | -0.021486 | 0.00046165 | -0.01 | 0.008 | 0.0007 | -0.0016 |  | 4 | 5.66 | 0 | 7.73 | 0 | -0.17 | 0 | -0.04 | 0.03 | 7.72 | 0.1707 | -0.021486 | 0.008 | -0.0016
5 | 5.66 | 0 | -7.73 | 0.17 | 0 | 0 | [6 empty] | 5.61 | -5.62 | -7.78 | 0.1723 |  |  |  |  | -0.05 | 0.13018 | 0.13945 | 0.039829 | -0.05 |  |  | 0.0023 | if COR>0: |  |  | 0.17 | -5.62 | [44 empty] | -0.05 | -0.033 | 0.13018 | 0.096175 | 0.10168 | 0.0076792 | -0.05 |  | 5 | 5.66 | 0 | -7.73 | 0.17 | 0 | 0 | 5.61 | -5.62 | -7.78 | 0.1723 | 0.097291 | 0.0032914 | 1.0833e-05 | -0.05 | -0.032 | 0.0023 | -3.4694e-18 |  | 5 | 5.66 | 0 | -7.73 | 0.17 | 0 | 0 | 5.61 | -5.62 | -7.78 | 0.1723 | 0.0032914 | -0.032 | -3.4694e-18
6 | 5.66 | 0 | -7.73 | -0.17 | 0 | 0 | [6 empty] | 5.62 | 5.65 | -7.73 | 0.1739 |  |  |  |  | -0.04 | -0.10018 | 0.10787 | 0.008248 | 0 |  |  | 0.0039 | Rx+=> displacement +Y |  |  | -0.17 | 5.65 | [44 empty] | -0.04 | -0.023 | -0.10018 | -0.13418 | 0.13613 | 0.042132 | 0 |  | 6 | 5.66 | 0 | -7.73 | -0.17 | 0 | 0 | 5.62 | 5.65 | -7.73 | 0.1739 | 0.13835 | 0.044351 | 0.001967 | 0 | 0.018 | 0.0039 | 0.0016 |  | 6 | 5.66 | 0 | -7.73 | -0.17 | 0 | 0 | 5.62 | 5.65 | -7.73 | 0.1739 | 0.044351 | 0.018 | 0.0016
7 | 5.66 | 0 | -7.73 | 0 | 0.17 | 0 | [6 empty] | 11.27 | -0.01 | -7.77 | 0.1676 |  |  |  |  | -0.14018 | -0.01 | 0.14053 | 0.040913 | -0.04 |  |  | 0.0024 | Ry+=> displacement -X | [48 empty] | -0.14018 | -0.12318 | -0.01 | -0.044 | 0.1308 | 0.036798 | -0.04 |  | 7 | 5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 11.27 | -0.01 | -7.77 | 0.1676 | 0.12718 | 0.033179 | 0.0011008 | -0.04 | -0.022 | 0.0024 | 0.0001 |  | 7 | 5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 11.27 | -0.01 | -7.77 | 0.1676 | 0.033179 | -0.022 | 0.0001
8 | 5.66 | 0 | -7.73 | 0 | -0.17 | 0 | [6 empty] | -0.03 | 0.02 | -7.75 | 0.1714 |  |  |  |  | 0.060175 | 0.02 | 0.063412 | -0.036206 | -0.02 |  |  | 0.0014 | [49 empty] | 0.060175 | 0.077175 | 0.02 | -0.014 | 0.078435 | -0.015565 | -0.02 |  | 8 | 5.66 | 0 | -7.73 | 0 | -0.17 | 0 | -0.03 | 0.02 | -7.75 | 0.1714 | 0.083915 | -0.010085 | 0.0001017 | -0.02 | -0.002 | 0.0014 | -0.0009 |  | 8 | 5.66 | 0 | -7.73 | 0 | -0.17 | 0 | -0.03 | 0.02 | -7.75 | 0.1714 | -0.010085 | -0.002 | -0.0009
zero | Commanded moves to position, from M2 hexapod EUI X: 0; Commanded moves to position, from M2 hexapod EUI Y: 0; Commanded moves to position, from M2 hexapod EUI Z: 0 | [70 empty] | zero | [6 empty] | 0 | 0 | 0 | [9 empty] | zero | [6 empty] | 0 | 0 | 0
9 | -5.66 | 0 | 7.73 | 0.17 | 0 | 0 | [6 empty] | -5.66 | -5.61 | 7.69 | 0.1716 |  |  |  |  | 0 | 0.14018 | 0.14018 | 0.040557 | -0.04 |  |  | 0.0016 |  |  |  | 0.17 | -5.61 | [44 empty] | 0 | 0.017 | 0.14018 | 0.10618 | 0.10753 | 0.013528 | -0.04 |  | 9 | -5.66 | 0 | 7.73 | 0.17 | 0 | 0 | -5.66 | -5.61 | 7.69 | 0.1716 | 0.10549 | 0.011495 | 0.00013213 | -0.04 | -0.022 | 0.0016 | -0.0007 |  | 9 | -5.66 | 0 | 7.73 | 0.17 | 0 | 0 | -5.66 | -5.61 | 7.69 | 0.1716 | 0.011495 | -0.022 | -0.0007
10 | -5.66 | 0 | 7.73 | -0.17 | 0 | 0 | [6 empty] | -5.65 | 5.7 | 7.76 | 0.1736 |  |  |  |  | 0.01 | -0.050175 | 0.051162 | -0.048456 | 0.03 |  |  | 0.0036 |  |  |  | -0.17 | 5.7 | [44 empty] | 0.01 | 0.027 | -0.050175 | -0.084175 | 0.088399 | -0.0056005 | 0.03 |  | 10 | -5.66 | 0 | 7.73 | -0.17 | 0 | 0 | -5.65 | 5.7 | 7.76 | 0.1736 | 0.092863 | -0.0011371 | 1.293e-06 | 0.03 | 0.048 | 0.0036 | 0.0013 |  | 10 | -5.66 | 0 | 7.73 | -0.17 | 0 | 0 | -5.65 | 5.7 | 7.76 | 0.1736 | -0.0011371 | 0.048 | 0.0013
11 | -5.66 | 0 | 7.73 | 0 | 0.17 | 0 | [6 empty] | 0.03 | 0.03 | 7.73 | 0.1688 |  |  |  |  | -0.060175 | 0.03 | 0.067239 | -0.032379 | 0 |  |  | 0.0012 | [49 empty] | -0.060175 | -0.043175 | 0.03 | -0.004 | 0.04336 | -0.05064 | 0 |  | 11 | -5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 0.03 | 0.03 | 7.73 | 0.1688 | 0.038812 | -0.055188 | 0.0030458 | 0 | 0.018 | 0.0012 | -0.0011 |  | 11 | -5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 0.03 | 0.03 | 7.73 | 0.1688 | -0.055188 | 0.018 | -0.0011
12 | -5.66 | 0 | 7.73 | 0 | -0.17 | 0 | [6 empty] | -11.31 | 0.05 | 7.72 | 0.1688 |  |  |  |  | 0.10018 | 0.05 | 0.11196 | 0.012342 | -0.01 |  |  | 0.0012 | Note: positioning error | [48 empty] | 0.10018 | 0.11718 | 0.05 | 0.016 | 0.11826 | 0.024263 | -0.01 |  | 12 | -5.66 | 0 | 7.73 | 0 | -0.17 | 0 | -11.31 | 0.05 | 7.72 | 0.1688 | 0.12286 | 0.028865 | 0.00083318 | -0.01 | 0.008 | 0.0012 | -0.0011 |  | 12 | -5.66 | 0 | 7.73 | 0 | -0.17 | 0 | -11.31 | 0.05 | 7.72 | 0.1688 | 0.028865 | 0.008 | -0.0011
13 | -5.66 | 0 | -7.73 | 0.17 | 0 | 0 | [6 empty] | -5.66 | -5.6 | -7.77 | 0.1728 |  |  |  |  | 0 | 0.15018 | 0.15018 | 0.050557 | -0.04 |  |  | 0.0028 |  | calculation can be |  | 0.17 | -5.6 | [44 empty] | 0 | 0.017 | 0.15018 | 0.11618 | 0.11741 | 0.023412 | -0.04 |  | 13 | -5.66 | 0 | -7.73 | 0.17 | 0 | 0 | -5.66 | -5.6 | -7.77 | 0.1728 | 0.11529 | 0.021294 | 0.00045342 | -0.04 | -0.022 | 0.0028 | 0.0005 |  | 13 | -5.66 | 0 | -7.73 | 0.17 | 0 | 0 | -5.66 | -5.6 | -7.77 | 0.1728 | 0.021294 | -0.022 | 0.0005
14 | -5.66 | 0 | -7.73 | -0.17 | 0 | 0 | [6 empty] | -5.64 | 5.68 | -7.72 | 0.173 |  |  |  |  | 0.02 | -0.070175 | 0.07297 | -0.026648 | 0.01 |  |  | 0.003 |  | improved |  | -0.17 | 5.68 | [44 empty] | 0.02 | 0.037 | -0.070175 | -0.10418 | 0.11055 | 0.016551 | 0.01 |  | 14 | -5.66 | 0 | -7.73 | -0.17 | 0 | 0 | -5.64 | 5.68 | -7.72 | 0.173 | 0.11511 | 0.021111 | 0.00044567 | 0.01 | 0.028 | 0.003 | 0.0007 |  | 14 | -5.66 | 0 | -7.73 | -0.17 | 0 | 0 | -5.64 | 5.68 | -7.72 | 0.173 | 0.021111 | 0.028 | 0.0007
15 | -5.66 | 0 | -7.73 | 0 | 0.17 | 0 | [6 empty] | 0.01 | 0.02 | -7.75 | 0.1687 |  |  |  |  | -0.080175 | 0.02 | 0.082632 | -0.016986 | -0.02 |  |  | 0.0013 | Go TO------> | start column CJ | [47 empty] | -0.080175 | -0.063175 | 0.02 | -0.014 | 0.064708 | -0.029292 | -0.02 |  | 15 | -5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 0.01 | 0.02 | -7.75 | 0.1687 | 0.060608 | -0.033392 | 0.001115 | -0.02 | -0.002 | 0.0013 | -0.001 |  | 15 | -5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 0.01 | 0.02 | -7.75 | 0.1687 | -0.033392 | -0.002 | -0.001
16 | -5.66 | 0 | -7.73 | 0 | -0.17 | 0 | [6 empty] | -11.35 | 0.05 | -7.75 | 0.1689 |  |  |  |  | 0.060175 | 0.05 | 0.078237 | -0.021381 | -0.02 |  |  | 0.0011 |  | ==============> |  | 0.17 | -5.6175 | [44 empty] | 0.060175 | 0.077175 | 0.05 | 0.016 | 0.078816 | -0.015184 | -0.02 |  | 16 | -5.66 | 0 | -7.73 | 0 | -0.17 | 0 | -11.35 | 0.05 | -7.75 | 0.1689 | 0.083197 | -0.010803 | 0.0001167 | -0.02 | -0.002 | 0.0011 | -0.0012 |  | 16 | -5.66 | 0 | -7.73 | 0 | -0.17 | 0 | -11.35 | 0.05 | -7.75 | 0.1689 | -0.010803 | -0.002 | -0.0012
17 | 0 | 5.66 | 7.73 | 0.17 | 0 | 0 | [6 empty] | -0.04 | 0.02 | 7.71 | 0.1722 |  |  |  |  | -0.04 | 0.11018 | 0.11721 | 0.017594 | -0.02 |  |  | 0.0022 |  |  |  | -0.17 | 5.6825 | [44 empty] | -0.04 | -0.023 | 0.11018 | 0.076175 | 0.079572 | -0.014428 | -0.02 |  | 17 | 0 | 5.66 | 7.73 | 0.17 | 0 | 0 | -0.04 | 0.02 | 7.71 | 0.1722 | 0.075357 | -0.018643 | 0.00034758 | -0.02 | -0.002 | 0.0022 | -0.0001 |  | 17 | 0 | 5.66 | 7.73 | 0.17 | 0 | 0 | -0.04 | 0.02 | 7.71 | 0.1722 | -0.018643 | -0.002 | -0.0001
18 | 0 | 5.66 | 7.73 | -0.17 | 0 | 0 | [6 empty] | -0.03 | 11.31 | 7.77 | 0.1755 |  |  |  |  | -0.03 | -0.10018 | 0.10457 | 0.0049529 | 0.04 |  |  | 0.0055 | [49 empty] | -0.03 | -0.013 | -0.10018 | -0.13418 | 0.1348 | 0.040804 | 0.04 |  | 18 | 0 | 5.66 | 7.73 | -0.17 | 0 | 0 | -0.03 | 11.31 | 7.77 | 0.1755 | 0.13741 | 0.043408 | 0.0018843 | 0.04 | 0.058 | 0.0055 | 0.0032 |  | 18 | 0 | 5.66 | 7.73 | -0.17 | 0 | 0 | -0.03 | 11.31 | 7.77 | 0.1755 | 0.043408 | 0.058 | 0.0032
19 | 0 | 5.66 | 7.73 | 0 | 0.17 | 0 | [6 empty] | 5.67 | 5.65 | 7.74 | 0.1673 |  |  |  |  | -0.080175 | -0.01 | 0.080796 | -0.018822 | 0.01 |  |  | 0.0027 | [49 empty] | -0.080175 | -0.063175 | -0.01 | -0.044 | 0.076988 | -0.017012 | 0.01 |  | 19 | 0 | 5.66 | 7.73 | 0 | 0.17 | 0 | 5.67 | 5.65 | 7.74 | 0.1673 | 0.074789 | -0.019211 | 0.00036907 | 0.01 | 0.028 | 0.0027 | 0.0004 |  | 19 | 0 | 5.66 | 7.73 | 0 | 0.17 | 0 | 5.67 | 5.65 | 7.74 | 0.1673 | -0.019211 | 0.028 | 0.0004
20 | 0 | 5.66 | 7.73 | 0 | -0.17 | 0 | [6 empty] | -5.71 | 5.68 | 7.74 | 0.1702 |  |  |  |  | 0.040175 | 0.02 | 0.044878 | -0.05474 | 0.01 |  |  | 0.0002 | [49 empty] | 0.040175 | 0.057175 | 0.02 | -0.014 | 0.058864 | -0.035136 | 0.01 |  | 20 | 0 | 5.66 | 7.73 | 0 | -0.17 | 0 | -5.71 | 5.68 | 7.74 | 0.1702 | 0.064457 | -0.029543 | 0.00087277 | 0.01 | 0.028 | 0.0002 | -0.0021 |  | 20 | 0 | 5.66 | 7.73 | 0 | -0.17 | 0 | -5.71 | 5.68 | 7.74 | 0.1702 | -0.029543 | 0.028 | -0.0021
21 | 0 | 5.66 | -7.73 | 0.17 | 0 | 0 | [6 empty] | -0.02 | -0.01 | -7.75 | 0.1707 |  |  |  |  | -0.02 | 0.080175 | 0.082632 | -0.016986 | -0.02 |  |  | 0.0007 | [49 empty] | -0.02 | -0.003 | 0.080175 | 0.046175 | 0.046273 | -0.047727 | -0.02 |  | 21 | 0 | 5.66 | -7.73 | 0.17 | 0 | 0 | -0.02 | -0.01 | -7.75 | 0.1707 | 0.043221 | -0.050779 | 0.0025785 | -0.02 | -0.002 | 0.0007 | -0.0016 |  | 21 | 0 | 5.66 | -7.73 | 0.17 | 0 | 0 | -0.02 | -0.01 | -7.75 | 0.1707 | -0.050779 | -0.002 | -0.0016
22 | 0 | 5.66 | -7.73 | -0.17 | 0 | 0 | [6 empty] | -0.02 | 11.32 | -7.71 | 0.1748 |  |  |  |  | -0.02 | -0.090175 | 0.092366 | -0.0072515 | 0.02 |  |  | 0.0048 | [49 empty] | -0.02 | -0.003 | -0.090175 | -0.12418 | 0.12421 | 0.030211 | 0.02 |  | 22 | 0 | 5.66 | -7.73 | -0.17 | 0 | 0 | -0.02 | 11.32 | -7.71 | 0.1748 | 0.12719 | 0.033191 | 0.0011016 | 0.02 | 0.038 | 0.0048 | 0.0025 |  | 22 | 0 | 5.66 | -7.73 | -0.17 | 0 | 0 | -0.02 | 11.32 | -7.71 | 0.1748 | 0.033191 | 0.038 | 0.0025
23 | 0 | 5.66 | -7.73 | 0 | 0.17 | 0 | [6 empty] | 5.62 | 5.66 | -7.73 | 0.168 |  |  |  |  | -0.13018 | 0 | 0.13018 | 0.030557 | 0 |  |  | 0.002 | [49 empty] | -0.13018 | -0.11318 | 0 | -0.034 | 0.11817 | 0.024172 | 0 |  | 23 | 0 | 5.66 | -7.73 | 0 | 0.17 | 0 | 5.62 | 5.66 | -7.73 | 0.168 | 0.11433 | 0.020328 | 0.00041322 | 0 | 0.018 | 0.002 | -0.0003 |  | 23 | 0 | 5.66 | -7.73 | 0 | 0.17 | 0 | 5.62 | 5.66 | -7.73 | 0.168 | 0.020328 | 0.018 | -0.0003
24 | 0 | 5.66 | -7.73 | 0 | -0.17 | 0 | [6 empty] | -5.68 | 5.67 | -7.73 | 0.1708 |  |  |  |  | 0.070175 | 0.01 | 0.070884 | -0.028734 | 0 |  |  | 0.0008 | [49 empty] | 0.070175 | 0.087175 | 0.01 | -0.024 | 0.090419 | -0.0035814 | 0 |  | 24 | 0 | 5.66 | -7.73 | 0 | -0.17 | 0 | -5.68 | 5.67 | -7.73 | 0.1708 | 0.096048 | 0.0020483 | 4.1954e-06 | 0 | 0.018 | 0.0008 | -0.0015 |  | 24 | 0 | 5.66 | -7.73 | 0 | -0.17 | 0 | -5.68 | 5.67 | -7.73 | 0.1708 | 0.0020483 | 0.018 | -0.0015
25 | 0 | -5.66 | 7.73 | 0.17 | 0 | 0 | [6 empty] | -0.04 | -11.23 | 7.65 | 0.176 |  |  |  |  | -0.04 | 0.18018 | 0.18456 | 0.084944 | -0.08 |  |  | 0.006 | [49 empty] | -0.04 | -0.023 | 0.18018 | 0.14618 | 0.14797 | 0.053974 | -0.08 |  | 25 | 0 | -5.66 | 7.73 | 0.17 | 0 | 0 | -0.04 | -11.23 | 7.65 | 0.176 | 0.1443 | 0.050302 | 0.0025303 | -0.08 | -0.062 | 0.006 | 0.0037 |  | 25 | 0 | -5.66 | 7.73 | 0.17 | 0 | 0 | -0.04 | -11.23 | 7.65 | 0.176 | 0.050302 | -0.062 | 0.0037
26 | 0 | -5.66 | 7.73 | -0.17 | 0 | 0 | [6 empty] | -0.03 | 0.09 | 7.72 | 0.1695 |  |  |  |  | -0.03 | -0.0001752 | 0.030001 | -0.069617 | -0.01 |  |  | 0.0005 |  |  |  | 12:50 Faulted after 5 minutes moving to this point Drive Fault and Simulink Fault. The faults were cleared with the Clear Error command | [45 empty] | -0.03 | -0.013 | -0.0001752 | -0.034175 | 0.036564 | -0.057436 | -0.01 |  | 26 | 0 | -5.66 | 7.73 | -0.17 | 0 | 0 | -0.03 | 0.09 | 7.72 | 0.1695 | 0.038026 | -0.055974 | 0.0031331 | -0.01 | 0.008 | 0.0005 | -0.0018 |  | 26 | 0 | -5.66 | 7.73 | -0.17 | 0 | 0 | -0.03 | 0.09 | 7.72 | 0.1695 | -0.055974 | 0.008 | -0.0018
27 | 0 | -5.66 | 7.73 | 0 | 0.17 | 0 | [6 empty] | 5.63 | -5.56 | 7.78 | 0.1665 |  |  |  |  | -0.12018 | 0.1 | 0.15634 | 0.056722 | 0.05 |  |  | 0.0035 | [49 empty] | -0.12018 | -0.10318 | 0.1 | 0.066 | 0.12248 | 0.028479 | 0.05 |  | 27 | 0 | -5.66 | 7.73 | 0 | 0.17 | 0 | 5.63 | -5.56 | 7.78 | 0.1665 | 0.11665 | 0.022651 | 0.00051305 | 0.05 | 0.068 | 0.0035 | 0.0012 |  | 27 | 0 | -5.66 | 7.73 | 0 | 0.17 | 0 | 5.63 | -5.56 | 7.78 | 0.1665 | 0.022651 | 0.068 | 0.0012
28 | 0 | -5.66 | 7.73 | 0 | -0.17 | 0 | [6 empty] | -5.72 | -5.55 | 7.68 | 0.1706 |  |  |  |  | 0.030175 | 0.11 | 0.11406 | 0.014446 | -0.05 |  |  | 0.0006 | [49 empty] | 0.030175 | 0.047175 | 0.11 | 0.076 | 0.089451 | -0.0045489 | -0.05 |  | 28 | 0 | -5.66 | 7.73 | 0 | -0.17 | 0 | -5.72 | -5.55 | 7.68 | 0.1706 | 0.089729 | -0.0042712 | 1.8243e-05 | -0.05 | -0.032 | 0.0006 | -0.0017 |  | 28 | 0 | -5.66 | 7.73 | 0 | -0.17 | 0 | -5.72 | -5.55 | 7.68 | 0.1706 | -0.0042712 | -0.032 | -0.0017
Back to zero | 0 | 0 | 0 | 0 | 0 | 0 | [6 empty] | -0.05 | 0.03 | -0.02 | [16 empty] | lunchtime | [51 empty] | -0.02 |  | Back to zero | 0 | 0 | 0 | 0 | 0 | 0 | -0.05 | 0.03 | -0.02 |  |  |  |  | -0.02 |  |  |  |  | Back to zero | 0 | 0 | 0 | 0 | 0 | 0 | -0.05 | 0.03 | -0.02
29 | 0 | -5.66 | -7.73 | 0.17 | 0 | 0 | [6 empty] | -0.04 | -11.22 | -7.82 | 0.1732 |  |  |  |  | 0.02 | 0.15018 | 0.1515 | 0.051883 | -0.09 |  |  | 0.0032 |  |  |  | 13:58 Tried to move position 29, Fault. Following Error, Drive Fault, Simulink Fault. Faults cleared with Clear error. | [45 empty] | 0.02 | 0.037 | 0.15018 | 0.11618 | 0.12192 | 0.027925 | -0.09 |  | 29 | 0 | -5.66 | -7.73 | 0.17 | 0 | 0 | -0.04 | -11.22 | -7.82 | 0.1732 | 0.12726 | 0.033264 | 0.0011065 | -0.09 | -0.072 | 0.0032 | 0.0009 |  | 29 | 0 | -5.66 | -7.73 | 0.17 | 0 | 0 | -0.04 | -11.22 | -7.82 | 0.1732 | 0.033264 | -0.072 | 0.0009
30 | 0 | -5.66 | -7.73 | -0.17 | 0 | 0 | [6 empty] | -0.04 | 0.07 | -7.76 | 0.1736 |  |  |  |  | 0.02 | -0.060175 | 0.063412 | -0.036206 | -0.03 |  |  | 0.0036 | [49 empty] | 0.02 | 0.037 | -0.060175 | -0.094175 | 0.10118 | 0.0071828 | -0.03 |  | 30 | 0 | -5.66 | -7.73 | -0.17 | 0 | 0 | -0.04 | 0.07 | -7.76 | 0.1736 | 0.092863 | -0.0011371 | 1.293e-06 | -0.03 | -0.012 | 0.0036 | 0.0013 |  | 30 | 0 | -5.66 | -7.73 | -0.17 | 0 | 0 | -0.04 | 0.07 | -7.76 | 0.1736 | -0.0011371 | -0.012 | 0.0013
31 | 0 | -5.66 | -7.73 | 0 | 0.17 | 0 | [6 empty] | 5.59 | -5.58 | -7.8 | 0.1671 |  |  |  |  | -0.10018 | 0.04 | 0.10787 | 0.008248 | -0.07 |  |  | 0.0029 | [49 empty] | -0.10018 | -0.083175 | 0.04 | 0.006 | 0.083391 | -0.010609 | -0.07 |  | 31 | 0 | -5.66 | -7.73 | 0 | 0.17 | 0 | 5.59 | -5.58 | -7.8 | 0.1671 | 0.098101 | 0.0041013 | 1.6821e-05 | -0.07 | -0.052 | 0.0029 | 0.0006 |  | 31 | 0 | -5.66 | -7.73 | 0 | 0.17 | 0 | 5.59 | -5.58 | -7.8 | 0.1671 | 0.0041013 | -0.052 | 0.0006
32 | 0 | -5.66 | -7.73 | 0 | -0.17 | 0 | [6 empty] | -5.71 | -5.57 | -7.79 | 0.1704 |  |  |  |  | 0.10018 | 0.05 | 0.11196 | 0.012342 | -0.06 |  |  | 0.0004 | [49 empty] | 0.10018 | 0.11718 | 0.05 | 0.016 | 0.11826 | 0.024263 | -0.06 |  | 32 | 0 | -5.66 | -7.73 | 0 | -0.17 | 0 | -5.71 | -5.57 | -7.79 | 0.1704 | 0.05438 | -0.03962 | 0.0015697 | -0.06 | -0.042 | 0.0004 | -0.0019 |  | 32 | 0 | -5.66 | -7.73 | 0 | -0.17 | 0 | -5.71 | -5.57 | -7.79 | 0.1704 | -0.03962 | -0.042 | -0.0019
zero | 0 | 0 | 0 | 0 | 0 | 0 | [6 empty] | -0.07 | 0.05 | -0.03 | 0.0011 | [60 empty] | Average | -0.013787 |  | 0.027307 |  | 0.08831 | 0.010271 | mm |  | zero | 0 | 0 | 0 | 0 | 0 | 0 | -0.07 | 0.05 | -0.03 | 0.0011 | [8 empty] | zero | 0 | 0 | 0 | 0 | 0 | 0 | -0.07 | 0.05 | -0.03 | 0.0011
[20 empty] | Average | -0.017419 | 0.031285 | 0.099618 | -2.4174e-17 | -0.017097 | mm |  | 0.0022812 | [48 empty] | Stdev | 0.065606 |  | 0.078962 |  | 0.045401 | 0.038074 | mm |  | Average | [10 empty] | 0.094045 | 4.5007e-05 |  | -0.017879 | -0.00024033 | 0.0022812 | -0.00058438 |  | Average | [10 empty] | 4.5007e-05 | 0.0001875 | -1.875e-05
[20 empty] | Stdev | 0.066198 | 0.077518 | 0.03823 | 0.03823 | 0.033486 | mm |  | 0.0015478 | [8 empty] | Absolute Accuracy Test (Moog 3.3.13) | [22 empty] | Traslational Range Z (Moog 3.3.6) (taken from absolute accuracy test data) | [25 empty] | STDEV | [10 empty] | 0.029303 | 0.029303 |  | 0.032669 | 0.031517 | 0.0015478 | 0.0016852 |  | STDEV | [10 empty] | 0.029303 | 0.031329 | 0.0016082
| Commanded moves to position, from M2 hexapod EUI X: 01/07/2021, Absolute Accuracy Test (Moog 3.3.13); Commanded moves to position, from M2 hexapod EUI X: EUI Actual Struts Position e.u. (um); Commanded moves to position, from M2 hexapod EUI X: laser tracker measured data | [12 empty] | Error (um) | [8 empty] | 2021-01-07 | Camera Hexapod Commanded position (um) | [6 empty] | EUI Actual Struts Position e.u. (um) |  |  |  |  | laser tracker data (mm) |  |  | Error (um) | [6 empty] | 2021-01-07 | Cam.Hexapod Commanded position (um) | [5 empty] | EUI Actual Struts Position e.u. (um) | [5 empty] | laser tracker data (mm) |  |  | Error (um) | [9 empty] | SUM | [12 empty] | 0.10008
Max range | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | spec: 25 times the resolution per axis | [11 empty] | Z | [8 empty] | Max range | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Z | [6 empty] | Max range | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Z | [9 empty] | RMS | [11 empty] | 0.055923
7.6 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | 0 | 0 | 0.61 | 0.07 | [12 empty] | spec<25 | crosstalkX | crosstalkY | [6 empty] | 7.6mm | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | 0 | 0 | 0.61 | 0.07 | spec<25 | [6 empty] | 7.6mm | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | 0 | 0 | 0.61 | 0.07 | spec<25
25% | 0 | 0 | 1.9 | 0 | 0 | 0 | -1555 | -1554 | -1555 | -1555 | -1555 | -1554 | 0.03 | 0.63 | 1.97 | [12 empty] | 0 | 0.03 | 0.02 |  | spec= | 0.025 |  |  |  | 25% | 0 | 0 | 1.9 | 0 | 0 | 0 | -1555 | -1554 | -1555 | -1555 | -1555 | -1554 | 0.03 | 0.63 | 1.97 | 0 | [6 empty] | 100% | 0 | 0 | 7.6 | 0 | 0 | 0 | -6202 | -6202 | -6203 | -6203 | -6203 | -6204 | 0 | 0.65 | 7.67 | 0
50% | 0 | 0 | 3.8 | 0 | 0 | 0 | -3106 | -3106 | -3107 | -3107 | -3107 | -3017 | 0 | 0.62 | 3.87 | [12 empty] | 4.4409e-13 | 0 | 0.01 | [6 empty] | 50% | 0 | 0 | 3.8 | 0 | 0 | 0 | -3106 | -3106 | -3107 | -3107 | -3107 | -3017 | 0 | 0.62 | 3.87 | 4.4409e-13 | [6 empty] | -100% | 0 | 0 | -7.6 | 0 | 0 | 0 | 6241 | 6241 | 6242 | 6242 | 6242 | 6242 | 0.02 | 0.62 | -7.53 | 0
75% | 0 | 0 | 5.7 | 0 | 0 | 0 | -4656 | -4655 | -4656 | -4656 | -4656 | -4656 | 0.03 | 0.63 | 5.77 | [12 empty] | 0 | 0.03 | 0.02 | [6 empty] | 75% | 0 | 0 | 5.7 | 0 | 0 | 0 | -4656 | -4655 | -4656 | -4656 | -4656 | -4656 | 0.03 | 0.63 | 5.77 | 0
100% | 0 | 0 | 7.6 | 0 | 0 | 0 | -6202 | -6202 | -6203 | -6203 | -6203 | -6204 | 0 | 0.65 | 7.67 | [12 empty] | 0 | 0 | 0.04 | [6 empty] | 100% | 0 | 0 | 7.6 | 0 | 0 | 0 | -6202 | -6202 | -6203 | -6203 | -6203 | -6204 | 0 | 0.65 | 7.67 | 0
-25% | 0 | 0 | -1.9 | 0 | 0 | 0 | 1557 | 1557 | 1557 | 1557 | 1557 | 1557 | 0.02 | 0.65 | -1.83 | [12 empty] | 2.2204e-13 | 0.02 | 0.04 | [6 empty] | -25% | 0 | 0 | -1.9 | 0 | 0 | 0 | 1557 | 1557 | 1557 | 1557 | 1557 | 1557 | 0.02 | 0.65 | -1.83 | 2.2204e-13
-50% | 0 | 0 | -3.8 | 0 | 0 | 0 | 3116 | 3116 | 3116 | 3116 | 3116 | 3116 | 0 | 0.61 | -3.73 | [12 empty] | 0 | 0 | 0 | [6 empty] | -50% | 0 | 0 | -3.8 | 0 | 0 | 0 | 3116 | 3116 | 3116 | 3116 | 3116 | 3116 | 0 | 0.61 | -3.73 | 0
-75% | 0 | 0 | -5.7 | 0 | 0 | 0 | 4677 | 4677 | 4678 | 4678 | 4678 | 4678 | 0 | 0.64 | -5.63 | [12 empty] | 8.8818e-13 | 0 | 0.03 | [6 empty] | -75% | 0 | 0 | -5.7 | 0 | 0 | 0 | 4677 | 4677 | 4678 | 4678 | 4678 | 4678 | 0 | 0.64 | -5.63 | 8.8818e-13
-100% | 0 | 0 | -7.6 | 0 | 0 | 0 | 6241 | 6241 | 6242 | 6242 | 6242 | 6242 | 0.02 | 0.62 | -7.53 | [12 empty] | 8.8818e-13 | 0.02 | 0.01 | [6 empty] | -100% | 0 | 0 | -7.6 | 0 | 0 | 0 | 6241 | 6241 | 6242 | 6242 | 6242 | 6242 | 0.02 | 0.62 | -7.53 | 8.8818e-13
Max range mm | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [5 empty] | EUI Actual Struts Position e.u. (um) 1: EUI Actual Struts Position. | [15 empty] | Y | crosstalkX | crosstalkZ | [6 empty] | Max range | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [5 empty] | EUI Actual Struts Position. | [8 empty] | Y
7.6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.01 | 0.62 | 0.07 | [12 empty] | spec<125 | [8 empty] | 7.6mm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.01 | 0.62 | 0.07 | spec<125
25% | 0 | 1.9 | 0 | 0 | 0 | 0 | 549 | 549 | -1088 | 549 | 549 | -1089 | 0.02 | 2.52 | 0.07 | [12 empty] | 0 | 0.03 | 0 |  | spec= | 0.125 |  |  |  | 25% | 0 | 1.9 | 0 | 0 | 0 | 0 | 549 | 549 | -1088 | 549 | 549 | -1089 | 0.02 | 2.52 | 0.07 | 0
50% | 0 | 3.8 | 0 | 0 | 0 | 0 | 1105 | 1105 | -2172 | 1105 | 1104 | -2172 | 0.02 | 4.39 | 0.08 | [12 empty] | 30 | 0.03 | 0.01 | [6 empty] | 50% | 0 | 3.8 | 0 | 0 | 0 | 0 | 1105 | 1105 | -2172 | 1105 | 1104 | -2172 | 0.02 | 4.39 | 0.08 | 30
75% | 0 | 5.7 | 0 | 0 | 0 | 0 | 1667 | 1667 | -3250 | 1666 | 1667 | -3251 | 0.02 | 6.3 | 0.09 | [12 empty] | 20 | 0.03 | 0.02 |  | 11:15 Faultr. Drive Fault, Simulink Fault. Clear error did not work to remove the fault. Exiting EUI and reloading EUI clear the error. |  |  |  |  | 75% | 0 | 5.7 | 0 | 0 | 0 | 0 | 1667 | 1667 | -3250 | 1666 | 1667 | -3251 | 0.02 | 6.3 | 0.09 | 20
100% | 0 | 7.6 | 0 | 0 | 0 | 0 | 2237 | 2237 | -4324 | 2235 | 2235 | -4324 | 0.04 | 8.19 | 0.1 | [12 empty] | 30 | 0.05 | 0.03 | [6 empty] | 100% | 0 | 7.6 | 0 | 0 | 0 | 0 | 2237 | 2237 | -4324 | 2235 | 2235 | -4324 | 0.04 | 8.19 | 0.1 | 30
-25% | 0 | -1.9 | 0 | 0 | 0 | 0 | -543 | -542 | 1093 | -542 | -542 | 1094 | 0.34 | -1.36 | 0.1 | [12 empty] | 80 | 0.35 | 0.03 | [6 empty] | -25% | 0 | -1.9 | 0 | 0 | 0 | 0 | -543 | -542 | 1093 | -542 | -542 | 1094 | 0.34 | -1.36 | 0.1 | 80
-50% | 0 | -3.8 | 0 | 0 | 0 | 0 | -1078 | -1078 | 2192 | -1078 | -1077 | 2192 | 0.31 | -3.25 | 0.09 | [12 empty] | 70 | 0.32 | 0.02 | [6 empty] | -50% | 0 | -3.8 | 0 | 0 | 0 | 0 | -1078 | -1078 | 2192 | -1078 | -1077 | 2192 | 0.31 | -3.25 | 0.09 | 70
-75% | 0 | -5.7 | 0 | 0 | 0 | 0 | -1607 | -1607 | 3294 | -1606 | -1606 | 3295 | -0.25 | -4.97 | 0.01 | [12 empty] | 110 | -0.24 | -0.06 | [6 empty] | -75% | 0 | -5.7 | 0 | 0 | 0 | 0 | -1607 | -1607 | 3294 | -1606 | -1606 | 3295 | -0.25 | -4.97 | 0.01 | 110
-100% | 0 | -7.6 | 0 | 0 | 0 | 0 | -2129 | -2129 | 4402 | -2128 | -2128 | 4402 | -0.29 | -6.86 | 0.01 | [12 empty] | 120 | -0.28 | -0.06 | [6 empty] | -100% | 0 | -7.6 | 0 | 0 | 0 | 0 | -2129 | -2129 | 4402 | -2128 | -2128 | 4402 | -0.29 | -6.86 | 0.01 | 120
Max range mm | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [5 empty] | EUI Actual Struts Position e.u. (um) 1: EUI Actual Struts Position. | [15 empty] | X | [8 empty] | Max range | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [5 empty] | EUI Actual Struts Position. | [8 empty] | X
7.6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.27 | 0.68 | 0.04 | [12 empty] | spec<125 | [8 empty] | 7.6mm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.27 | 0.68 | 0.04 | spec<125
25% | 1.9 | 0 | 0 | 0 | 0 | 0 | -942 | 948 | 3 | -942 | 947 | 3 | 1.61 | 0.66 | 0.04 | [12 empty] | 20 | [8 empty] | 25% | 1.9 | 0 | 0 | 0 | 0 | 0 | -942 | 948 | 3 | -942 | 947 | 3 | 1.61 | 0.66 | 0.04 | 20
50% | 3.8 | 0 | 0 | 0 | 0 | 0 | -1879 | 1901 | 14 | -1879 | 1901 | 14 | 3.51 | 0.65 | 0.04 | [12 empty] | 20 |  |  |  | spec= | 0.125 |  |  |  | 50% | 3.8 | 0 | 0 | 0 | 0 | 0 | -1879 | 1901 | 14 | -1879 | 1901 | 14 | 3.51 | 0.65 | 0.04 | 20
75% | 5.7 | 0 | 0 | 0 | 0 | 0 | -2810 | 2860 | 33 | -2810 | 2859 | 33 | 5.41 | 0.65 | 0.04 | [12 empty] | 20 | [8 empty] | 75% | 5.7 | 0 | 0 | 0 | 0 | 0 | -2810 | 2860 | 33 | -2810 | 2859 | 33 | 5.41 | 0.65 | 0.04 | 20
100% | 7.6 | 0 | 0 | 0 | 0 | 0 | -3736 | 3824 | 58 | -3735 | 3823 | 58 | 7.3 | 0.65 | 0.04 | [12 empty] | 30 | [8 empty] | 100% | 7.6 | 0 | 0 | 0 | 0 | 0 | -3736 | 3824 | 58 | -3735 | 3823 | 58 | 7.3 | 0.65 | 0.04 | 30
-25% | -1.9 | 0 | 0 | 0 | 0 | 0 | 948 | -942 | 4 | 948 | -943 | 4 | -2.16 | 0.68 | 0.03 | [12 empty] | 10 | [5 empty] | 5.48 | -0.22 |  | -25% | -1.9 | 0 | 0 | 0 | 0 | 0 | 948 | -942 | 4 | 948 | -943 | 4 | -2.16 | 0.68 | 0.03 | 10
-50% | -3.8 | 0 | 0 | 0 | 0 | 0 | 1901 | -1879 | 14 | 1900 | -1879 | 14 | -4.04 | 0.69 | 0.03 | [12 empty] | 30 | [8 empty] | -50% | -3.8 | 0 | 0 | 0 | 0 | 0 | 1901 | -1879 | 14 | 1900 | -1879 | 14 | -4.04 | 0.69 | 0.03 | 30
-75% | -5.7 | 0 | 0 | 0 | 0 | 0 | 2860 | -2810 | 32 | 2859 | -2810 | 33 | -5.94 | 0.7 | 0.03 | [12 empty] | 30 | [8 empty] | -75% | -5.7 | 0 | 0 | 0 | 0 | 0 | 2860 | -2810 | 32 | 2859 | -2810 | 33 | -5.94 | 0.7 | 0.03 | 30
-100% | -7.6 | 0 | 0 | 0 | 0 | 0 | 3824 | -3735 | 58 | 3823 | -3735 | 58 | -7.82 | 0.7 | 0.03 | [12 empty] | 50 | [8 empty] | -100% | -7.6 | 0 | 0 | 0 | 0 | 0 | 3824 | -3735 | 58 | 3823 | -3735 | 58 | -7.82 | 0.7 | 0.03 | 50
zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | -0.27 | 0.68 | 0.04 | [21 empty] | zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | -0.27 | 0.68 | 0.04
X and Y Translational Range (MOOG 3.3.4), COR= -2758400um | [19 empty] | calc displ | [6 empty] | expected | Error L.T. |  | Cross talk check | [6 empty] | X and Y Translational Range (MOOG 3.3.4), COR= -2758400um | [15 empty] | calc displ | expected | Error L.T. |  | Cross talk check
2021-02-07 | Commanded moves to position, from M2 hexapod EUI | [5 empty] | EUI Actual Struts Position e.u. (um) 2: EUI Actual Struts Position e.u. (um) | laser tracker data mm | [6 empty] | (XY data) | (x2+y2)1/2 | [5 empty] | XY | (meas.-set) | laser tracker data mm | [7 empty] | 2021-02-07 | Commanded moves to position, from M2 hexapod EUI | [6 empty] | EUI Actual Struts Position e.u. (um) |  |  |  |  | laser tracker data mm |  |  | (XY data) | XY | (meas.-set) | laser tracker data mm
Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z |  |  |  |  | mm | [6 empty] | mm | um | X | Y | Z | [5 empty] | Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | mm | mm | um | X | Y | Z
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | 0 | -0.26 | 0.69 | 0.03 | [12 empty] | spec<125 | [8 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | 0 | -0.26 | 0.69 | 0.03 |  |  | spec<125
1 | 7.6 | 0 | 0 | 0 | 0 | 0 | -3736 | 3824 | 58 | -3735 | 3823 | 58 | 7.31 | 0.67 | 0.03 |  |  |  |  | 7.57 | [6 empty] | 7.6 | 29.9736 | 7.57 | 0.02 | 0 | [5 empty] | 1 | 7.6 | 0 | 0 | 0 | 0 | 0 | -3736 | 3824 | 58 | -3735 | 3823 | 58 | 7.31 | 0.67 | 0.03 | 7.57 | 7.6 | 29.9736 | 7.57 | 0.02 | 0
2 | 5.374 | 5.374 | 0 | 0 | 0 | 0 | -1072 | 4256 | -3036 | -1072 | 4255 | -3037 | 5.1 | 6.02 | 0.05 |  | 7.56 |  |  | 7.559 | [6 empty] | 7.6 | 40.9824 | 5.36 | 5.33 | 0.02 | [5 empty] | 2 | 5.374 | 5.374 | 0 | 0 | 0 | 0 | -1072 | 4256 | -3036 | -1072 | 4255 | -3037 | 5.1 | 6.02 | 0.05 | 7.559 | 7.6 | 40.9824 | 5.36 | 5.33 | 0.02
3 | 0 | 7.6 | zcdsv | 0 | 0 | 0 | 2237 | 2236 | -4323 | 2235 | 2235 | -4323 | -0.23 | 8.25 | 0.06 |  |  |  |  | 7.5601 | [6 empty] | 7.6 | 39.9405 | 0.03 | 7.56 | 0.01 | [5 empty] | 3 | 0 | 7.6 | 0 | 0 | 0 | 0 | 2237 | 2236 | -4323 | 2235 | 2235 | -4323 | -0.23 | 8.25 | 0.06 | 7.5601 | 7.6 | 39.9405 | 0.03 | 7.56 | 0.01
4 | -5.374 | 5.374 | 0 | 0 | 0 | 0 | 4256 | -1072 | -3036 | 4255 | -1072 | -3036 | -5.61 | 6.03 | 0.05 |  |  |  |  | 7.559 | [6 empty] | 7.6 | 41.0089 | -5.35 | 5.34 | 0.01 | [5 empty] | 4 | -5.374 | 5.374 | 0 | 0 | 0 | 0 | 4256 | -1072 | -3036 | 4255 | -1072 | -3036 | -5.61 | 6.03 | 0.05 | 7.559 | 7.6 | 41.0089 | -5.35 | 5.34 | 0.01
5 | -7.6 | 0 | 0 | 0 | 0 | 0 | 3824 | -3736 | 59 | 3823 | -3735 | 59 | -7.82 | 0.73 | 0.02 |  |  |  |  | 7.5601 | [6 empty] | 7.6 | 39.8942 | -7.56 | 0.04 | 0.03 | [5 empty] | 5 | -7.6 | 0 | 0 | 0 | 0 | 0 | 3824 | -3736 | 59 | 3823 | -3735 | 59 | -7.82 | 0.73 | 0.02 | 7.5601 | 7.6 | 39.8942 | -7.56 | 0.04 | 0.03
6 | -5.374 | -5.374 | 0 | 0 | 0 | 0 | 1186 | -4175 | 3134 | 1186 | -4173 | 3134 | -5.63 | -4.66 | 0.01 |  |  |  |  | 7.5802 | [6 empty] | 7.6 | 19.7858 | -5.37 | -5.35 | 0.01 | [5 empty] | 6 | -5.374 | -5.374 | 0 | 0 | 0 | 0 | 1186 | -4175 | 3134 | 1186 | -4173 | 3134 | -5.63 | -4.66 | 0.01 | 7.5802 | 7.6 | 19.7858 | -5.37 | -5.35 | 0.01
7 | 0 | -7.6 | 0 | 0 | 0 | 0 | -2129 | -2129 | 4402 | -2128 | -2127 | 4402 | -0.28 | -6.82 | 0 |  | -7.51 |  |  | 7.51 | [6 empty] | 7.6 | 89.9734 | 0.02 | -7.51 | 0.01 | [5 empty] | 7 | 0 | -7.6 | 0 | 0 | 0 | 0 | -2129 | -2129 | 4402 | -2128 | -2127 | 4402 | -0.28 | -6.82 | 0 | 7.51 | 7.6 | 89.9734 | 0.02 | -7.51 | 0.01
8 | 5.374 | -5.374 | 0 | 0 | 0 | 0 | -4175 | 1187 | 3134 | -4174 | 1187 | 3134 | 5.08 | -4.62 | 0 |  |  |  |  | 7.5307 | [6 empty] | 7.6 | 69.2666 | 5.34 | -5.31 | 0 | [5 empty] | 8 | 5.374 | -5.374 | 0 | 0 | 0 | 0 | -4175 | 1187 | 3134 | -4174 | 1187 | 3134 | 5.08 | -4.62 | 0 | 7.5307 | 7.6 | 69.2666 | 5.34 | -5.31 | 0
[56 empty] | 0.0070711 | 0.014142 | 0.0099103
X and Y rotational Range (MOOG 3.3.8) | [15 empty] | plane-to-plane |  |  |  | calc Angl |  | expected | Error L.T. |  | Cross talk check |  |  |  |  | angl error | [6 empty] | X and Y rotational Range (MOOG 3.3.8) | [15 empty] | plane-to-plane | calc Angl | expected | Error L.T. |  | Cross talk check
2021-02-07 | COR= -2758400um | [5 empty] | EUI Actual Struts Position e.u. (um) 2: EUI Actual Struts Position e.u. (um) | laser tracker data mm |  |  | ¦Rot.ang.¦ |  |  |  | (RxRy) | (Rx2+Ry2)1/2 |  | (meas.-set) | laser tracker data mm |  |  |  | comments |  | calc.from | [6 empty] | 2021-02-07 | COR= -2758400um | [6 empty] | EUI Actual Struts Position e.u. (um) |  |  |  |  | laser tracker data mm |  |  | ¦Rot.ang.¦ | (RxRy) | RxRy | (meas.-set) | laser tracker data mm
Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | (Deg) |  |  |  | Deg |  | Deg | Deg | X | Y | Z |  |  |  | x and y | [6 empty] | Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | (Deg) | Deg | Deg | Deg | X | Y | Z
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | -0.27 | 0.72 | 0.02 | 0 | [6 empty] | spec<0.0021 |  |  |  | with SMR or Gauges on Rotator |  |  | displ. | [6 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | -0.27 | 0.72 | 0.02 | 0 |  |  | spec<0.0021
1 | 0 | 0 | 0 | 0.24 | 0 | 0 | -4842 | -4842 | 5885 | -778 | -778 | 5885 | -0.26 | -7.26 | -0.04 | 0.2428 |  |  |  | 0.23856 |  | 0.24 | 0.0014373 | 0.01 | -7.98 | -0.06 | if COR<0: |  |  | 0.0014373 |  | 0.24 | -7.98 |  |  |  | 1 | 0 | 0 | 0 | 0.24 | 0 | 0 | -4842 | -4842 | 5885 | -778 | -778 | 5885 | -0.26 | -7.26 | -0.04 | 0.2428 | 0.23856 | 0.24 | 0.0014373 | 0.01 | -7.98 | -0.06
2 | 0 | 0 | 0 | 0.17 | 0.17 | 0 | -6165 | -637 | 2572 | -4944 | 3864 | 5845 | 5.41 | -4.93 | -0.02 | 0.2402 |  |  |  | 0.23942 |  | 0.24042 | 0.00099976 | 5.68 | -5.65 | -0.04 | Rx+=> displacement -Y |  |  | 0.00099976 |  | 0.17 | -5.6175 |  |  |  | 2 | 0 | 0 | 0 | 0.17 | 0.17 | 0 | -6165 | -637 | 2572 | -4944 | 3864 | 5845 | 5.41 | -4.93 | -0.02 | 0.2402 | 0.23942 | 0.24042 | 0.00099976 | 5.68 | -5.65 | -0.04
3 | 0 | 0 | 0 | 0 | 0.24 | 0 | -3780 | 3963 | -2224 | -6134 | 6269 | 2439 | 7.73 | 0.72 | 0 | 0.2375 |  |  |  | 0.23886 |  | 0.24 | 0.0011387 | 8 | 0 | -0.02 | Ry+=> displacement +X |  |  | 0.0011387 |  | -0.17 | 5.6825 |  |  |  | 3 | 0 | 0 | 0 | 0 | 0.24 | 0 | -3780 | 3963 | -2224 | -6134 | 6269 | 2439 | 7.73 | 0.72 | 0 | 0.2375 | 0.23886 | 0.24 | 0.0011387 | 8 | 0 | -0.02
4 | 0 | 0 | 0 | -0.17 | 0.17 | 0 | 859 | 6301 | -5697 | -3664 | 5101 | -2363 | 5.42 | 6.38 | 0.03 | 0.2414 |  |  |  | 0.23942 |  | 0.24042 | 0.00099864 | 5.69 | 5.66 | 0.01 | if COR>0: |  |  | 0.00099864 | [6 empty] | 4 | 0 | 0 | 0 | -0.17 | 0.17 | 0 | 859 | 6301 | -5697 | -3664 | 5101 | -2363 | 5.42 | 6.38 | 0.03 | 0.2414 | 0.23942 | 0.24042 | 0.00099864 | 5.69 | 5.66 | 0.01
5 | 0 | 0 | 0 | -0.24 | 0 | 0 | 5014 | 5014 | -5744 | 998 | 998 | -5745 | -0.27 | 8.71 | 0.04 | 0.2443 |  |  |  | 0.23826 |  | 0.24 | 0.001736 | 0 | 7.99 | 0.02 | Rx+=> displacement +Y |  |  | 0.001736 |  | -0.24 | 7.99 |  |  |  | 5 | 0 | 0 | 0 | -0.24 | 0 | 0 | 5014 | 5014 | -5744 | 998 | 998 | -5745 | -0.27 | 8.71 | 0.04 | 0.2443 | 0.23826 | 0.24 | 0.001736 | 0 | 7.99 | 0.02
6 | 0 | 0 | 0 | -0.17 | -0.17 | 0 | 6301 | 858 | -2362 | 5101 | -3664 | -5698 | -5.93 | 6.4 | 0.04 | 0.2409 |  |  |  | 0.23921 |  | 0.24042 | 0.0012112 | -5.66 | 5.68 | 0.02 | Ry+=> displacement -X |  |  | 0.0012112 | [6 empty] | 6 | 0 | 0 | 0 | -0.17 | -0.17 | 0 | 6301 | 858 | -2362 | 5101 | -3664 | -5698 | -5.93 | 6.4 | 0.04 | 0.2409 | 0.23921 | 0.24042 | 0.0012112 | -5.66 | 5.68 | 0.02
7 | 0 | 0 | 0 | 0 | -0.24 | 0 | 3963 | -3780 | 2439 | 6269 | -6135 | -2224 | -8.27 | 0.77 | 0 | 0.2395 |  |  |  | 0.23886 |  | 0.24 | 0.0011359 | -8 | 0.05 | -0.02 |  |  |  | 0.0011359 | [5 empty] | 14:21 Fault State, Drive fault, Simulink Fault. Clear error commando did not worked | 7 | 0 | 0 | 0 | 0 | -0.24 | 0 | 3963 | -3780 | 2439 | 6269 | -6135 | -2224 | -8.27 | 0.77 | 0 | 0.2395 | 0.23886 | 0.24 | 0.0011359 | -8 | 0.05 | -0.02
8 | 0 | 0 | 0 | 0.17 | -0.17 | 0 | -638 | -6165 | 5844 | 3864 | -4944 | 2572 | -5.95 | -4.88 | -0.03 | 0.2436 |  |  |  | 0.23837 |  | 0.24042 | 0.0020512 | -5.68 | -5.6 | -0.05 |  |  |  | 0.0020512 | [6 empty] | 8 | 0 | 0 | 0 | 0.17 | -0.17 | 0 | -638 | -6165 | 5844 | 3864 | -4944 | 2572 | -5.95 | -4.88 | -0.03 | 0.2436 | 0.23837 | 0.24042 | 0.0020512 | -5.68 | -5.6 | -0.05
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.27 | 0.74 | 0.01 | [8 empty] | 0.0070711 | 0.035355 | 0.03151 | [10 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.27 | 0.74 | 0.01 |  |  |  |  | 0.0070711 | 0.035355 | 0.03151
|  |  |  | 0.24042 | 8.1179 | [7 empty] | -0.27 | 0.73 | 0.015 | [6 empty] | averages= | 0.0013386 | 0.005 | 0.025 | -0.0175 |  |  |  | 0.0013386 | [5 empty] | deg |  |  |  |  | 0.24042 | 8.1179 | [7 empty] | -0.27 | 0.73 | 0.015 |  | overall cross-talk, mm RMS= |  |  | 0.01291 | 0.022174 | 0.027011
|  |  |  | 0.24042 | 5.7502 | [16 empty] | stdev= | 0.00037855 | 6.0585 | 6.0368 | 0.03151 |  |  |  | 0.00037855 | [5 empty] | deg |  |  |  |  | angle Rx | offset Y | [7 empty] | av. zero offsets X,Y,Z
[13 empty] | 0.01 | -7.98 | -0.06 | [38 empty] | overall cross-talk, mm RMS= |  |  | 0.01291 | 0.022174 | 0.027011
[13 empty] | 5.68 | -5.65 | -0.04
[13 empty] | 8 | 0 | -0.02
[13 empty] | 5.69 | 5.66 | 0.01
[13 empty] | 0 | 7.99 | 0.02
[13 empty] | -5.66 | 5.68 | 0.02
[13 empty] | -8 | 0.05 | -0.02
[13 empty] | -5.68 | -5.6 | -0.05
From X and Y rotational range above | [29 empty] | 0.00205 | 0.123
Moog 3.3.13 Rotation Absolute Accuracy tests (assembled from positioning and maximum range data | [26 empty] | expected | Error L.T. | [8 empty] | Moog 3.3.13 Rotation Absolute Accuracy tests (assembled from positioning and maximum range data) | [17 empty] | expected | Error L.T.
Rx rotational absolute accuracy test | [15 empty] | plane-to-plane |  | SMR |  | calc Angl | [6 empty] | XY | (meas.-set) | [8 empty] | Rx rotational absolute accuracy test | [15 empty] | plane-to-plane | calc Angl | XY | (meas.-set)
2021-02-07 | COR= -2758400um | [5 empty] | EUI Actual Struts Position e.u. (um) 2: EUI Actual Struts Position e.u. (um) | laser tracker data mm |  |  | ¦Rot.ang.¦ |  | Offset |  | Deg | [6 empty] | mm | Deg | [8 empty] | 2021-02-07 | COR= -2758400um | [6 empty] | EUI Actual Struts Position e.u. (um) |  |  |  |  | laser tracker data mm |  |  | ¦Rot.ang.¦ | Deg | mm | Deg
Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | (Deg) |  | 19.05 mm |  | Rx | [7 empty] | spec<0.0021 | [8 empty] | Move# | X | Y | Z | Rx | Ry | Rz | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | (Deg) | Rx |  | spec<0.0021
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -0.01 | 0.005 | 0 |  |  |  | 0.00029277 | 0.00029277 | [5 empty] | 0 | 0.00029277 | [8 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -0.01 | 0.005 | 0 | 0.00029277 | 0 | 0.00029277
1 | 0 | 0 | 0 | 0.24 | 0 | 0 | -4842 | -4842 | 5885 | -778 | -778 | 5885 | 0.01 | -7.99 | -0.055 | 0.2428 |  |  |  | 0.23856 | 0.23856 | [5 empty] | 0.24 | 0.0014375 |  |  |  | if COR<0: |  |  |  |  | 1 | 0 | 0 | 0 | 0.24 | 0 | 0 | -4842 | -4842 | 5885 | -778 | -778 | 5885 | 0.01 | -7.99 | -0.055 | 0.2428 | 0.23856 | 0.24 | 0.0014375
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.005 | 0.0003 |  |  |  | -0.00029858 | 0.00029858 | [5 empty] | 0 | 0.00029858 |  |  |  | Rx+=> displacement -Y |  |  |  |  | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.005 | 0.0003 | -0.00029858 | 0 | 0.00029858
5 | 0 | 0 | 0 | -0.24 | 0 | 0 | 5014 | 5014 | -5744 | 998 | 998 | -5745 | 0 | 7.98 | 0.025 | 0.2443 |  |  |  | -0.23826 | 0.23826 | [5 empty] | -0.24 | 0.001736 |  |  |  | Ry+=> displacement +X |  |  |  |  | 5 | 0 | 0 | 0 | -0.24 | 0 | 0 | 5014 | 5014 | -5744 | 998 | 998 | -5745 | 0 | 7.98 | 0.025 | 0.2443 | -0.23826 | -0.24 | 0.001736
Ry rotational absolute accuracy test | [19 empty] | Ry | [16 empty] | Ry rotational absolute accuracy test | [16 empty] | Ry
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -0.01 | 0.005 | 0 |  |  |  | 0 | 0 | [5 empty] | 0 | 0 | [8 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -0.01 | 0.005 | 0 | 0 | 0 | 0
3 | 0 | 0 | 0 | 0 | 0.24 | 0 | -3780 | 3963 | -2224 | -6134 | 6269 | 2439 | 8 | -0.01 | -0.015 | 0.2375 |  |  |  | 0.23775 | 0.23775 | [5 empty] | 0.24 | 0.0022539 | [8 empty] | 3 | 0 | 0 | 0 | 0 | 0.24 | 0 | -3780 | 3963 | -2224 | -6134 | 6269 | 2439 | 8 | -0.01 | -0.015 | 0.2375 | 0.23775 | 0.24 | 0.0022539
pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.005 | 0.0003 |  |  |  | 0 | 0 | [5 empty] | 0 | 0 | [8 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.005 | 0.0003 | 0 | 0 | 0
7 | 0 | 0 | 0 | 0 | -0.24 | 0 | 3963 | -3780 | 2439 | 6269 | -6135 | -2224 | -8 | 0.04 | -0.015 | 0.2395 |  |  |  | -0.23775 | 0.23775 | [5 empty] | -0.24 | 0.0022539 | [8 empty] | 7 | 0 | 0 | 0 | 0 | -0.24 | 0 | 3963 | -3780 | 2439 | 6269 | -6135 | -2224 | -8 | 0.04 | -0.015 | 0.2395 | -0.23775 | -0.24 | 0.0022539
[16 empty] | 0.001025 |  |  |  | -0.23993 | -0.0019203
Rotator tests (see last worksheet)
| 05/07/2021, 3.4.1 Axis of Rotation | [5 empty] | 1: Rotator Position by EUI |  |  |  | angl. error
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | -0.28 | 0.72 | 0.05 | 0.0141
[6 empty] | 90 | EUI Actual Struts Position e.u. (um) 6: 90 | 0 | 0 | 0 |  | 0 | 0 | 89.975
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0035
[6 empty] | -90 | EUI Actual Struts Position e.u. (um) 6: -90 | 0 | 0 | 0 |  | 0 | 0 | -90.0926
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0031
| 05/07/2021, 3.4.2 Rotator Absolute Accuracy | [5 empty] | 1: Rotator Position by EUI
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0031
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 29.9944
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 59.9857
[6 empty] | 90 | EUI Actual Struts Position e.u. (um) 6: 90 | 0 | 0 | 0 |  | 0 | 0 | 89.9745
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 59.9851
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 29.994
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0042
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.0246
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0573
[6 empty] | -90 | EUI Actual Struts Position e.u. (um) 6: -90 | 0 | 0 | 0 |  | 0 | 0 | -90.0738
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0576
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.0257
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0003
MOOG 3.4.5.3
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | 0
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 30.0113
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 60.0024 | [12 empty] | 14:38 Fault, Safety Interlock Fault, Simulnk Fault Declared. After send the Clear command Faulted again. This repeated 4 times in total. Stoped EUI and restarted again
[6 empty] | 90 | EUI Actual Struts Position e.u. (um) 6: 90 | 0 | 0 | 0 |  | 0 | 0 | 89.8251
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 60.0024
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 30.0108
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | 0.0093 | [12 empty] | 14:56 EUI stop automatically after moving to 30 degree. (it moved untl 30.005 deg. Then EUI stop) 14:59 Fault Again.Fault State. Safety Interlock. Simulink Fault declared
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.0165
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0492
[6 empty] | -90 | EUI Actual Struts Position e.u. (um) 6: -90 | 0 | 0 | 0 |  | 0 | 0 | -90.0952
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0491
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.017
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | 0.0088
[7 empty] | 1: LASER TRACKER SMRx | LASER TRACKER SMR-x
[7 empty] | 1: X; EUI Actual Struts Position e.u. (um) 2: Y; EUI Actual Struts Position e.u. (um) 4: (W/R to frame 4 SMR center) | X | Y | (W/R to frame 4 SMR center)
[6 empty] | 0 | 1: 759.66; EUI Actual Struts Position e.u. (um) 2: 0.07; EUI Actual Struts Position e.u. (um) 6: 0 | -760.38 | 0.55
[6 empty] | 30 | 1: 657.93; EUI Actual Struts Position e.u. (um) 2: 379.97; EUI Actual Struts Position e.u. (um) 6: 30 | -658.77 | -379.38
[6 empty] | 60 | 1: 379.95; EUI Actual Struts Position e.u. (um) 2: 658.13; EUI Actual Struts Position e.u. (um) 6: 60 | -380.85 | -657.64
[6 empty] | 90 | 1: 0.08; EUI Actual Struts Position e.u. (um) 2: 760.22; EUI Actual Struts Position e.u. (um) 6: 90 | -0.94 | -759.66
[6 empty] | 60 | 1: 379.93; EUI Actual Struts Position e.u. (um) 2: 658.13; EUI Actual Struts Position e.u. (um) 6: 60 | -380.85 | -657.65
[6 empty] | 30 | 1: 657.92; EUI Actual Struts Position e.u. (um) 2: 379.97; EUI Actual Struts Position e.u. (um) 6: 30 | -658.77 | -379.4
[6 empty] | 0 | 1: 759.64; EUI Actual Struts Position e.u. (um) 2: 0.07; EUI Actual Struts Position e.u. (um) 6: 0 | -760.38 | 0.53
[6 empty] | -30 | 1: 657.71; EUI Actual Struts Position e.u. (um) 2: -380.01; EUI Actual Struts Position e.u. (um) 6: -30 | -658.35 | 380.57
[6 empty] | -60 | 1: 379.24; EUI Actual Struts Position e.u. (um) 2: -658.08; EUI Actual Struts Position e.u. (um) 6: -60 | -379.94 | 658.67
[6 empty] | -90 | 1: -1.11; EUI Actual Struts Position e.u. (um) 2: -759.68; EUI Actual Struts Position e.u. (um) 6: -90 | 0.39 | 760.18
[6 empty] | -60 | 1: 379.25; EUI Actual Struts Position e.u. (um) 2: -658.07; EUI Actual Struts Position e.u. (um) 6: -60 | -379.94 | 658.66
[6 empty] | -30 | 1: 657.72; EUI Actual Struts Position e.u. (um) 2: -380.02; EUI Actual Struts Position e.u. (um) 6: -30 | -658.35 | 380.56
[6 empty] | 0 | 759.64 | 0.07 |  | new center coord. |  | 0 | -760.39 | 0.52
[7 empty] | 494.4277 | 0.059231 |  | x | y |  | -495.1938 | 0.50077
[10 empty] | -0.38308 | 0.28
MOOG 3.4.5.4 - 3.4.5.5 - 3.4.5.6
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0.38 | 0.05 |  | 0 | 0 | -0.0022
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 29.9895
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 59.983
[6 empty] | 90 | EUI Actual Struts Position e.u. (um) 6: 90 | 0 | 0 | 0 |  | 0 | 0 | 89.9837
[6 empty] | 60 | EUI Actual Struts Position e.u. (um) 6: 60 | 0 | 0 | 0 |  | 0 | 0 | 59.9838
[6 empty] | 30 | EUI Actual Struts Position e.u. (um) 6: 30 | 0 | 0 | 0 |  | 0 | 0 | 29.99
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0022
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.0088
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0182 | -60.0178
[6 empty] | -90 | EUI Actual Struts Position e.u. (um) 6: -90 | 0 | 0 | 0 |  | 0 | 0 | -90.0419
[6 empty] | -60 | EUI Actual Struts Position e.u. (um) 6: -60 | 0 | 0 | 0 |  | 0 | 0 | -60.0179
[6 empty] | -30 | EUI Actual Struts Position e.u. (um) 6: -30 | 0 | 0 | 0 |  | 0 | 0 | -30.0097
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0017
3.4.5.4 - 3.4.5.5 - 3.4.5.6
[6 empty] | 0 | EUI Actual Struts Position e.u. (um) 6: 0 | 0 | 0 | 0 |  | 0 | 0 | -0.0023
30
60
90
60
30
0
-30
-60
-90
-60
-30
0

## Pos.3.3.1_Accuracy_Range_300322
| Cam. Hexapod re-verification With ComCam mounted on Rotator on Integrator at El 0Deg | [12 empty] | Commanded Positions |  |  |  |  | pos # |  | measured position 2022 |  |  | (mm and Deg) |  |  |  | Positioning absolurte accuracy specs | [9 empty] | absolurte accuracy 3.3.13 March 2022 | [5 empty] | Zero position of the 3 SMRs
[13 empty] | COR= | -2758400um | [5 empty] | (data from big X,Y,Z table left plus plane-to-plane | [5 empty] | Z: 25um | = | 0.025 | mm |  | XY: 125um | = | 0.125 | mm | [9 empty] | p1 | 859.0765 | 0 | -19.05
30/03/2022 |  |  |  |  | note: SMRs 1.5" placed at +X, +Y and -X locations on Rotator | [7 empty] | pivot = | 1938mm | (front of rotator surface at L1S1) |  |  |  |  | angles obtained directly from the LT SA software) | [5 empty] | RxRy: 205e-5Deg |  | = | 0.00205 | Deg; Rz: 1500e-5Deg = |  |  | 0.015 | Deg |  | LTS-206 3.3.13. Translation absolute accuracy | [7 empty] | p2 | -0.28566 | 859.0765 | -19.05
| raw LT data; coordinate values in mm |  |  |  |  | Includes X,Y&Z shifts due to tilt |  |  | removed displacement due to tilt | [19 empty] | Measured Error | [7 empty] | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [6 empty] | p3 | -859.0761 | 0.92068 | -19.05 |  |  |  | error
|  | X | Y | Z |  | X | Y | Z | X | Y | Z |  | X | Y | Z | Rx | Ry | Rz |  | X | Y | Z | Rx | Ry | Rz | x | y | XY |  | Z | Rx | Ry | RxRy | Rz |  | Max range | X | Y | Z | Rx | Ry | Rz | [5 empty] | X | Y | Z | mm
p1... X+ | p1 | 859.0765 | 0 | -19.05 |  | 0 | 0 | 0 |  |  |  | zero | 0 | 0 | 0 | 0 | 0 | 0 | zero | [16 empty] | 7.6 | 0 | 0 | 0 | 0 | 0 | 0 |  | -0.2807 | 859.0457 | -19.0643 |  | 0.004995 | -0.030307 | -0.014265 | -0.014265 | bar: XY
zero | p2 | -0.28566 | 859.0765 | -19.05 |  | 0 | 0 | 0 |  |  |  | 1 | 5.66 | 0 | 7.73 | 0.17 | 0 | 0 | 1 | 5.7144 | 0.011833 | 7.6782 | 0.1708 | 0.0016092 | -0.0043485 | 0.05443 | 0.011833 | 0.052966 | -0.051794 | -0.023794 | 0.0008 | 0.0016092 | 0.0023859 | -0.0043485 |  | 25% | 0 | 0 | 1.9 | 0 | 0 | 0 |  | -0.23852 | 859.0666 | -17.137 |  | 0.047177 | -0.009363 | 1.913 | 0.012984
0 | p3 | -859.0761 | 0.92068 | -19.05 |  | 0 | 0 | 0 |  |  |  | 2 | 5.66 | 0 | 7.73 | -0.17 | 0 | 0 | 2 | 5.6966 | 0.011907 | 7.7194 | -0.1688 | 0.0010833 | -0.0014813 | 0.036643 | 0.011907 | 0.039313 | -0.010563 | 0.017437 | 0.0012 | 0.0010833 | 0.0020446 | -0.0014813 |  | 50% | 0 | 0 | 3.8 | 0 | 0 | 0 |  | -0.28165 | 859.0596 | -15.2447 |  | 0.004046 | -0.016412 | 3.8053 | 0.005261 |  | absolute accuracy cross talk
| p1 | 864.7821 | -5.7712 | -11.3143 |  | 5.7056 | -5.7712 | 7.7357 | 5.7056 | -0.021062 | 7.7357 | 3 | 5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 3 | 5.6838 | 0.0077377 | 7.6738 | 0.00022876 | 0.1676 | -0.0060606 | 0.023843 | 0.0077377 | 0.028347 | -0.056175 | -0.028175 | 0.00022876 | -0.0024 | 0.0017002 | -0.0060606 |  | 75% | 0 | 0 | 5.7 | 0 | 0 | 0 |  | -0.24034 | 859.075 | -13.3424 |  | 0.04536 | -0.000961 | 5.7076 | 0.007572
1 | p2 | 5.4587 | 853.3373 | -8.7967 |  | 5.7443 | -5.7392 | 10.2533 | 5.7443 | 0.010977 | 7.6478 | 4 | 5.66 | 0 | 7.73 | 0 | -0.17 | 0 | 4 | 5.6723 | 0.002012 | 7.7146 | 5.9483e-05 | -0.1711 | -0.0028578 | 0.012315 | 0.002012 | 0.019152 | -0.015378 | 0.012622 | 5.9483e-05 | -0.0011 | 0.00042396 | -0.0028578 |  | 100% | 0 | 0 | 7.6 | 0 | 0 | 0 |  | -0.26216 | 859.0643 | -11.4339 |  | 0.023539 | -0.011726 | 7.6161 | 0.016056 |  |  | 24/03/2022 |  | (data from the measurements of 3.3.4 and 3.3.8)
| p3 | -853.3827 | -4.7839 | -11.399 |  | 5.6934 | -5.7046 | 7.6511 | 5.6934 | 0.045584 | 7.6511 | 5 | 5.66 | 0 | -7.73 | 0.17 | 0 | 0 | 5 | 5.6644 | -0.040673 | -7.8164 | 0.1696 | 0.00013126 | -0.00012514 | 0.0044397 | -0.040673 | 0.024317 | -0.086423 | -0.058423 | -0.0004 | 0.00013126 | 0.0010252 | -0.00012514 |  | -25% | 0 | 0 | -1.9 | 0 | 0 | 0 |  | -0.24444 | 859.0586 | -20.9395 |  | 0.041257 | -0.017432 | -1.8895 | 0.010479 |  | X | Y | Z |  | Rx |  | Ry |  | Rz
| p4 | 864.7701 | 5.758 | -11.3043 |  | 5.6937 | 5.758 | 7.7457 | 5.6937 | 0.0077758 | 7.7457 | 6 | 5.66 | 0 | -7.73 | -0.17 | 0 | 0 | 6 | 5.6582 | -0.039277 | -7.778 | -0.1708 | -5.2841e-05 | -0.00054136 | -0.0017873 | -0.039277 | 0.025203 | -0.047986 | -0.019986 | -0.0008 | -5.2841e-05 | 0.0011911 | -0.00054136 |  | -50% | 0 | 0 | -3.8 | 0 | 0 | 0 |  | -0.28222 | 859.0397 | -22.8743 |  | 0.003478 | -0.036314 | -3.8243 | -0.024291 |  | 0.0086265 | -0.012763 | -0.003359 | 0.00022876 | 0.00052876 | 0.0016092 | 0.0013092 | -0.0043485 | -0.0031485
2 | p5 | 5.4296 | 864.8241 | -13.8846 |  | 5.7153 | 5.7476 | 5.1654 | 5.7153 | -0.0025322 | 7.7709 | 7 | 5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 7 | 5.652 | -0.045467 | -7.8164 | -0.0013442 | 0.169 | -0.0036488 | -0.0080335 | -0.045467 | 0.033699 | -0.086352 | -0.058352 | -0.0013442 | -0.001 | 0.0015731 | -0.0036488 |  | -75% | 0 | 0 | -5.7 | 0 | 0 | 0 |  | -0.26721 | 859.0618 | -24.7429 |  | 0.018487 | -0.014216 | -5.6929 | 0.007054 |  | 0.049528 | -0.008105 | 0.009791 | 5.9483e-05 | 0.00035948 | 0.0010833 | 0.00078332 | -0.0014813 | -0.00028132
| p6 | -853.3951 | 6.7013 | -11.4083 |  | 5.681 | 5.7807 | 7.6417 | 5.681 | 0.030479 | 7.6417 | 8 | 5.66 | 0 | -7.73 | 0 | -0.17 | 0 | 8 | 5.6638 | -0.024571 | -7.7819 | -0.00072641 | -0.171 | -0.0013117 | 0.0037785 | -0.024571 | 0.0097991 | -0.051923 | -0.023923 | -0.00072641 | -0.001 | 0.00097378 | -0.0013117 |  | -100% | 0 | 0 | -7.6 | 0 | 0 | 0 |  | -0.26459 | 859.0557 | -26.6726 |  | 0.02111 | -0.020291 | -7.6226 | -0.022586 |  | 0.0042405 | -0.016814 | -0.013049 | -0.0013442 | -0.0010442 | 0.00013126 | -0.00016874 | -0.0060606 | -0.0048606
| p7 | 870.5007 | -0.035345 | -13.8907 |  | 11.4242 | -0.035345 | 5.1593 | 5.674 | -0.035345 | 7.7648 | 9 | -5.66 | 0 | 7.73 | 0.17 | 0 | 0 | 9 | -5.6299 | -0.0035578 | 7.6883 | 0.1697 | 0.00088924 | 0.0002482 | 0.030078 | -0.0035578 | 0.024162 | -0.041717 | -0.013717 | -0.0003 | 0.00088924 | 0.001666 | 0.0002482 |  | Max range mm | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [15 empty] | 0.039879 | -0.011047 | 0.004742 | -0.00072641 | -0.00042641 | -5.2841e-05 | -0.00035284 | -0.0028578 | -0.0016578
3 | p8 | 11.1783 | 859.0775 | -11.3739 |  | 11.4639 | 0.001016 | 7.6761 | 5.7138 | 0.001016 | 7.6761 | 10 | -5.66 | 0 | 7.73 | -0.17 | 0 | 0 | 10 | -5.6327 | -0.0044259 | 7.7357 | -0.1698 | 0.00080777 | 0.0022221 | 0.027322 | -0.0044259 | 0.021405 | 0.0057386 | 0.033739 | 0.0002 | 0.00080777 | 0.0015078 | 0.0022221 |  | 7.6 | 0 | 0 | 0 | 0 | 0 | 0 |  | -0.26844 | 859.0256 | -19.0727 |  | 0.012258 | -0.020439 | -0.008699 | -0.020439 |  | 0.023683 | -0.0072335 | -0.0035815 | -0.00011892 | 0.00018108 | 0.00088924 | 0.00058924 | -0.00012514 | 0.0010749
| p9 | -847.6622 | 0.97822 | -8.864 |  | 11.4139 | 0.057542 | 10.186 | 5.6637 | 0.057542 | 7.5806 | 11 | -5.66 | 0 | 7.73 | 0 | 0.17 | 0 | 11 | -5.6177 | -0.0040223 | 7.7276 | -0.00011892 | 0.1691 | -0.0010929 | 0.042296 | -0.0040223 | 0.034806 | -0.0024107 | 0.025589 | -0.00011892 | -0.0009 | 0.00037644 | -0.0010929 |  | 25% | 0 | 1.9 | 0 | 0 | 0 | 0 |  | -0.22882 | 860.9518 | -19.0357 |  | 0.051878 | 1.9058 | 0.028319 | 0.005848 |  | 0.04186 | -0.013554 | 0.010692 | -0.00031966 | -1.9659e-05 | 0.00080777 | 0.00050777 | -0.00054136 | 0.00065864
| p10 | 858.9924 | -0.018033 | -8.7778 |  | -0.084039 | -0.018033 | 10.2722 | 5.6661 | -0.018033 | 7.6667 | 12 | -5.66 | 0 | 7.73 | 0 | -0.17 | 0 | 12 | -5.6116 | -0.010812 | 7.7313 | -0.00031966 | -0.1711 | 0.00032154 | 0.048365 | -0.010812 | 0.038861 | 0.0012863 | 0.029286 | -0.00031966 | -0.0011 | 0.00065578 | 0.00032154 |  | 50% | 0 | 3.8 | 0 | 0 | 0 | 0 |  | -0.27626 | 862.8076 | -19.079 |  | 0.004435 | 3.7616 | -0.014954 | -0.038387 |  | 0.004438 | -0.017293 | 0.0034125 | -0.0011517 | -0.00085167 | 0.00075103 | 0.00045103 | -0.0036488 | -0.0024488
4 | p11 | -0.34881 | 859.0748 | -11.3201 |  | -0.063153 | -0.001698 | 7.7299 | 5.687 | -0.001698 | 7.7299 | 13 | -5.66 | 0 | -7.73 | 0.17 | 0 | 0 | 13 | -5.6346 | -0.025363 | -7.7879 | 0.1708 | 0.00075103 | -0.00075152 | 0.025403 | -0.025363 | 0.017527 | -0.0579 | -0.0299 | 0.0008 | 0.00075103 | 0.0015702 | -0.00075152 |  | 75% | 0 | 5.7 | 0 | 0 | 0 | 0 |  | -0.2463 | 864.7257 | -19.0414 |  | 0.034397 | 5.6797 | 0.022627 | -0.020273 |  | 0.035415 | -0.005771 | 0.0066 | 4.1489e-06 | 0.00030415 | 0.001091 | 0.00079099 | -0.0013117 | -0.00011167
| p12 | -859.1624 | 0.94644 | -13.9082 |  | -0.08639 | 0.025767 | 5.1418 | 5.6638 | 0.025767 | 7.7472 | 14 | -5.66 | 0 | -7.73 | -0.17 | 0 | 0 | 14 | -5.6231 | -0.035343 | -7.7379 | -0.16905 | 0.001091 | -0.00099307 | 0.036902 | -0.035343 | 0.03256 | -0.0079344 | 0.020066 | 0.00095 | 0.001091 | 0.0019417 | -0.00099307 |  | 100% | 0 | 7.6 | 0 | 0 | 0 | 0 |  | -0.25687 | 866.606 | -19.0713 |  | 0.023827 | 7.56 | -0.007282 | -0.040008 |  | 0.011177 | -0.015563 | 0.001187 | -0.0025565 | -0.0022565 | -0.0009784 | -0.0012784 | 0.0002482 | 0.0014482
| p13 | 864.7405 | -5.7907 | -26.8174 |  | 5.664 | -5.7907 | -7.7674 | 5.664 | -0.040562 | -7.7674 | 15 | -5.66 | 0 | -7.73 | 0 | 0.17 | 0 | 15 | -5.6207 | -0.038955 | -7.7629 | -0.0011517 | 0.1679 | -0.0022453 | 0.039274 | -0.038955 | 0.036592 | -0.032887 | -0.004887 | -0.0011517 | -0.0021 | 0.001946 | -0.0022453 |  | -25% | 0 | -1.9 | 0 | 0 | 0 | 0 |  | -0.23824 | 857.1719 | -19.039 |  | 0.042463 | -1.8741 | 0.024976 | 0.025854 |  |  |  |  | -0.0040142 | -0.0037142 | -0.0014157 | -0.0017157 | 0.0022221 | 0.0034221
5 | p14 | 5.3912 | 853.2835 | -24.3036 |  | 5.6769 | -5.793 | -5.2536 | 5.6769 | -0.042814 | -7.8591 | 16 | -5.66 | 0 | -7.73 | 0 | -0.17 | 0 | 16 | -5.6577 | 0.00014033 | -7.7525 | 4.1489e-06 | -0.171 | 0.0029153 | 0.0022749 | 0.00014033 | 0.018801 | -0.022513 | 0.005487 | 4.1489e-06 | -0.001 | 0.00035827 | 0.0029153 |  | -50% | 0 | -3.8 | 0 | 0 | 0 | 0 |  | -0.2753 | 855.2513 | -19.0621 |  | 0.005398 | -3.7947 | 0.001874 | 0.005307 |  |  |  |  | -0.0033267 | -0.0030267 | -0.0014149 | -0.0017149 | -0.0010929 | 0.0001071
| p15 | -853.4236 | -4.8681 | -26.8728 |  | 5.6524 | -5.7888 | -7.8228 | 5.6524 | -0.038644 | -7.8228 | 17 | 0 | 5.66 | 7.73 | 0.17 | 0 | 0 | 17 | -0.033094 | 5.5842 | 7.7031 | 0.1717 | -0.0009784 | -0.0011233 | -0.033094 | -0.07579 | 0.072893 | -0.02686 | 0.0011397 | 0.0017 | -0.0009784 | 0.0015256 | -0.0011233 |  | -75% | 0 | -5.7 | 0 | 0 | 0 | 0 |  | -0.22836 | 853.3678 | -19.0571 |  | 0.052342 | -5.6782 | 0.006891 | 0.021779 |  |  |  |  | 0.0022417 | 0.0025417 | 0.00094834 | 0.00064834 | 0.00032154 | 0.0015215
| p16 | 864.7294 | 5.7095 | -26.8215 |  | 5.6529 | 5.7095 | -7.7715 | 5.6529 | -0.040643 | -7.7715 | 18 | 0 | 5.66 | 7.73 | -0.17 | 0 | 0 | 18 | -0.047885 | 5.5609 | 7.7345 | -0.1699 | -0.0014157 | -9.3957e-05 | -0.047885 | -0.099089 | 0.10045 | 0.004489 | 0.032489 | 0.0001 | -0.0014157 | 0.00072264 | -9.3957e-05 |  | -100% | 0 | -7.6 | 0 | 0 | 0 | 0 |  | -0.27946 | 851.4585 | -19.063 |  | 0.001244 | -7.5875 | 0.000972 | 0.012512 |  |  |  |  | 0.0013651 | 0.0016651 | 0.00065271 | 0.00035271 | -0.00075152 | 0.00044848
6 | p17 | 5.3861 | 864.7818 | -29.403 |  | 5.6718 | 5.7053 | -10.353 | 5.6718 | -0.044842 | -7.7475 | 19 | 0 | 5.66 | 7.73 | 0 | 0.17 | 0 | 19 | [16 empty] | Max range mm | 01/07/2021, Absolute Accuracy Test (Moog 3.3.13) | [18 empty] | 0.0020831 | 0.0023831 | 0.00037348 | 7.3476e-05 | -0.00099307 | 0.00020693 | bar: XY
| p18 | -853.4261 | 6.6385 | -26.865 |  | 5.6499 | 5.7178 | -7.8149 | 5.6499 | -0.032346 | -7.8149 | 20 | 0 | 5.66 | 7.73 | 0 | -0.17 | 0 | 20 | -0.069248 | 5.5735 | 7.7202 | -0.0025565 | -0.1708 | -0.00033237 | -0.069248 | -0.086473 | 0.10539 | -0.0097803 | 0.01822 | -0.0025565 | -0.0008 | 0.0027583 | -0.00033237 |  | 7.6 | 0 | 0 | 0 | 0 | 0 | 0 |  | -0.27095 | 859.0308 | -19.071 |  | -0.002553 | 0.004781 | 0.001981 | -0.002553 |  |  |  |  | 0.0026902 | 0.0029902 |  |  | -0.0022453 | -0.0010453
| p19 | 870.4811 | -0.073032 | -29.3986 |  | 11.4047 | -0.073032 | -10.3485 | 5.6545 | -0.073032 | -7.7431 | 21 | 0 | 5.66 | -7.73 | 0.17 | 0 | 0 | 21 | -0.04786 | 5.5481 | -7.7882 | 0.172 | -0.0014149 | -0.0017801 | -0.04786 | -0.11188 | 0.11113 | -0.05822 | -0.03022 | 0.002 | -0.0014149 | 0.0019368 | -0.0017801 |  | 25% | 1.9 | 0 | 0 | 0 | 0 | 0 |  | 1.6149 | 859.0392 | -19.0727 |  | 1.8956 | -0.006847 | -0.008737 | -0.004437 | [8 empty] | 0.0029153 | 0.0041153
7 | p20 | 11.1303 | 859.0302 | -26.8702 |  | 11.416 | -0.046261 | -7.8202 | 5.6658 | -0.046261 | -7.8202 | 22 | 0 | 5.66 | -7.73 | -0.17 | 0 | 0 | 22 | 0.032077 | 5.5817 | -7.7281 | -0.169 | 0.00094834 | -0.0036199 | 0.032077 | -0.078321 | 0.065174 | 0.0018893 | 0.029889 | 0.001 | 0.00094834 | 0.0018322 | -0.0036199 |  | 50% | 3.8 | 0 | 0 | 0 | 0 | 0 |  | 3.5164 | 859.0288 | -19.0751 |  | 3.7971 | -0.017216 | -0.011143 | -0.00286 | [8 empty] | -0.0011233 | 7.6698e-05
| p21 | -847.6903 | 0.90357 | -24.3303 |  | 11.3858 | -0.017109 | -5.2803 | 5.6356 | -0.017109 | -7.8858 | 23 | 0 | 5.66 | -7.73 | 0 | 0.17 | 0 | 23 | -0.0025925 | 5.5242 | -7.7808 | -0.0040142 | 0.169 | -0.0040025 | -0.0025925 | -0.13578 | 0.11945 | -0.05083 | -0.02283 | -0.0040142 | -0.001 | 0.0042249 | -0.0040025 |  | 75% | 5.7 | 0 | 0 | 0 | 0 | 0 |  | 5.4338 | 859.0249 | -19.0771 |  | 5.7145 | -0.021132 | -0.013143 | 0.014491 | [8 empty] | -9.3957e-05 | 0.001106
| p22 | 858.993 | -0.038368 | -24.269 |  | -0.083423 | -0.038368 | -5.219 | 5.6668 | -0.038368 | -7.8245 | 24 | 0 | 5.66 | -7.73 | 0 | -0.17 | 0 | 24 | 0.0072445 | 5.5475 | -7.7636 | -0.0033267 | -0.1718 | -0.001001 | 0.0072445 | -0.11252 | 0.095564 | -0.033611 | -0.005611 | -0.0033267 | -0.0018 | 0.0036943 | -0.001001 |  | 100% | 7.6 | 0 | 0 | 0 | 0 | 0 |  | 7.3344 | 859.0433 | -19.0639 |  | 7.6151 | -0.002741 | 0.000119 | 0.015071 | [8 empty] | -0.00033237 | 0.00086763
8 | p23 | -0.3742 | 859.0594 | -26.8312 |  | -0.088539 | -0.017079 | -7.7812 | 5.6616 | -0.017079 | -7.7812 | 25 | 0 | -5.66 | 7.73 | 0.17 | 0 | 0 | 25 | [16 empty] | -25% | -1.9 | 0 | 0 | 0 | 0 | 0 |  | -2.1666 | 859.0363 | -19.0676 |  | -1.8859 | -0.009677 | -0.003593 | 0.014141 | [8 empty] | -0.0017801 | -0.00058008 | bar: XY
| p24 | -859.1633 | 0.90241 | -29.3956 |  | -0.087228 | -0.018265 | -10.3456 | 5.6629 | -0.018265 | -7.7402 | 26 | 0 | -5.66 | 7.73 | -0.17 | 0 | 0 | 26 | 0.022078 | -5.5793 | 7.7474 | -0.1683 | 0.00065271 | -0.002025 | 0.022078 | 0.080664 | 0.098408 | 0.017384 | 0.045384 | 0.0017 | 0.00065271 | 0.0020199 | -0.002025 |  | -50% | -3.8 | 0 | 0 | 0 | 0 | 0 |  | -4.0576 | 859.0477 | -19.059 |  | -3.7769 | 0.001729 | 0.004951 | 0.023096 | [8 empty] | -0.0036199 | -0.0024199
| p25 | 859.0913 | -0.013646 | -19.0753 |  | 0.014831 | -0.013646 | -0.025308 |  |  |  | 27 | 0 | -5.66 | 7.73 | 0 | 0.17 | 0 | 27 | 0.025564 | -5.5842 | 7.7085 | 0.0022417 | 0.1668 | -0.002344 | 0.025564 | 0.075824 | 0.09412 | -0.021469 | 0.0065313 | 0.0022417 | -0.0032 | 0.0032278 | -0.002344 |  | -75% | -5.7 | 0 | 0 | 0 | 0 | 0 |  | -5.9425 | 859.0487 | -19.0577 |  | -5.6618 | 0.002674 | 0.006309 | 0.038185 | [8 empty] | -0.0040025 | -0.0028025
zero | p26 | -0.27654 | 859.0591 | -19.0885 |  | 0.009118 | -0.017418 | -0.038467 |  |  |  | 28 | 0 | -5.66 | 7.73 | 0 | -0.17 | 0 | 28 | 0.0082855 | -5.6138 | 7.6889 | 0.0013651 | -0.1732 | -0.00071831 | 0.0082855 | 0.046175 | 0.063199 | -0.041082 | -0.013082 | 0.0013651 | -0.0032 | 0.0027582 | -0.00071831 |  | -100% | -7.6 | 0 | 0 | 0 | 0 | 0 |  | -7.8296 | 859.0352 | -19.0626 |  | -7.5489 | -0.010834 | 0.001402 | 0.051078 | [8 empty] | -0.001001 | 0.00019897
0.018039 | p27 | -859.0823 | 0.91526 | -19.0834 |  | -0.006227 | -0.005416 | -0.033399 |  |  |  | 29 | 0 | -5.66 | -7.73 | 0.17 | 0 | 0 | 29 | [16 empty] | zero | 0 | 0 | 0 | 0 | 0 | 0 |  | -0.28566 | 859.0765 | -19.05 | [13 empty] | -0.002025 | -0.00082502
| p28 | 853.4596 | -5.7634 | -11.3558 |  | -5.6317 | -5.7497 | 7.7195 | -5.6317 | 0.0004622 | 7.7195 | 30 | 0 | -5.66 | -7.73 | -0.17 | 0 | 0 | 30 | 0.012633 | -5.5731 | -7.7285 | -0.169 | 0.00037348 | 0.0019082 | 0.012633 | 0.086941 | 0.10397 | 0.001549 | 0.029549 | 0.001 | 0.00037348 | 0.0013388 | 0.0019082 | [25 empty] | -0.002344 | -0.001144
9 | p29 | -5.911 | 853.3011 | -8.8308 |  | -5.6345 | -5.758 | 10.2577 | -5.6345 | -0.0077938 | 7.6522 | 31 | 0 | -5.66 | -7.73 | 0 | 0.17 | 0 | 31 | -0.0039259 | -5.5895 | -7.7695 | 0.0020831 | 0.168 | -0.00060687 | -0.0039259 | 0.070462 | 0.088563 | -0.039466 | -0.011466 | 0.0020831 | -0.002 | 0.0022883 | -0.00060687 |  | LTS-206 3.3.13. Rotation Absolute Accuracy tests (assembled from positioning and maximum range data) | [23 empty] | -0.00071831 | 0.00048169
| p30 | -864.7059 | -4.8383 | -11.3903 |  | -5.6236 | -5.7535 | 7.6931 | -5.6236 | -0.0033418 | 7.6931 | 32 | 0 | -5.66 | -7.73 | 0 | -0.17 | 0 | 32 | 0.01309 | -5.569 | -7.7294 | 0.0026902 | -0.1719 | 0.002008 | 0.01309 | 0.090995 | 0.10804 | 0.000633 | 0.028633 | 0.0026902 | -0.0019 | 0.0027643 | 0.002008 |  | Rx rotational absolute accuracy test | [8 empty] | Zero Values | [14 empty] | 0.0019082 | 0.0031082
| p31 | 853.4582 | 5.7515 | -11.3388 |  | -5.6331 | 5.7652 | 7.7366 | -5.6331 | 0.014991 | 7.7366 | [7 empty] | Aver | x: 0.010135; y: -0.016815; XY: 0.058064 | -0.028286 | -0.00028638 | 0.00017463 | -0.0006905 | 0.001808 | -0.0011546 |  | 23/3/2022 | [7 empty] | p1 | 859.0775 | -0.042503 | -19.0742 | [12 empty] | -0.00060687 | 0.00059313
10 | p32 | -5.9184 | 864.8 | -13.9113 |  | -5.6419 | 5.741 | 5.1772 | -5.6419 | -0.0092032 | 7.7826 | [7 empty] | Stdev | XY: 0.036056 | 0.027805 | 0.027805 |  |  | 0.00095712 | 0.0020433 |  |  | COR= -2758400um | [6 empty] | p2 | -859.0758 | 0.88609 | -19.0754 | [12 empty] | 0.002008 | 0.003208
| p33 | -864.7053 | 6.6464 | -11.3954 |  | -5.623 | 5.7311 | 7.688 | -5.623 | -0.019065 | 7.688 | [26 empty] | trans. in mm |  |  | rot. in Deg |  |  | reduced laser tracker data mm |  |  | measured rotation angle in Deg |  |  |  | error
| p34 | 859.2102 | -0.030021 | -13.8951 |  | 0.11891 | -0.016375 | 5.1803 | -5.6313 | -0.016375 | 7.7857 |  | the tilt induced displacement calculations: | [5 empty] | Notes: pos.# 19 faulted (last trajectory aborted)…continued to following positions, didn't repeat. | [16 empty] | Move# | X | Y | Z | Rx | Ry | Rz |  | X | Y | Z |  | Rx | Ry | Rz | Deg | max. | 0.049528 | -0.017293 | -0.013049 |  | -0.0037142 |  | -0.0017157 |  | -0.0048606
11 | p35 | -0.13787 | 859.063 | -11.3447 |  | 0.13868 | 0.003933 | 7.7438 | -5.6115 | 0.003933 | 7.7438 |  | Pivot | d= |  | 1938 | mm |  |  | X: pos.#s 25 & 29 operational errors (sent to wrong position, made wrong measurement sequence). | [11 empty] | 0 |  |  |  | 0.0001395 | -5e-07 | -0.001435 |  | 1.4782e-08 |  |  | 1.4782e-08 | aver. | 0.024316 | -0.012016 | 0.0018261 | -0.00032571 | -2.5715e-05 | 0.00031967 | 1.9674e-05 | -0.0011546 | 4.5443e-05
| p36 | -858.9425 | 0.91563 | -8.8247 |  | 0.13982 | 0.000375 | 10.2587 | -5.6104 | 0.000375 | 7.6533 |  |  | angle= |  | 0.17 | Deg |  |  | X: XY&Rz accuracy meets specs; Z positioning accuracy marginally meets specs, RxRy accuracy meets specs most of the time | [7 empty] | 25% |  |  |  | 0.06 |  |  |  | 0.0075875 | -2.0164 | -0.00146 |  | 0.059614 |  |  | -0.00038622 | stdev. | 0.01777 | 0.0042496 | 0.0075302 | 0.0019717 | 0.0019717 | 0.0009599 | 0.0009599 | 0.0020433 | 0.0020433
| p37 | 847.7249 | -0.027813 | -8.7943 |  | -11.3664 | -0.014167 | 10.281 | -5.6162 | -0.014167 | 7.6755 |  | offset displacement = |  |  | 5.7502 | mm |  | Data reduction method [where 3 SMRs were used (one at X+ one at Y+ and one at X-)]: | [16 empty] | 50 |  |  |  | 0.12 |  |  |  | 0.031523 | -4.0415 | -0.011987 |  | 0.11948 |  |  | -0.00051561
12 | p38 | -11.6373 | 859.0599 | -11.3345 |  | -11.3607 | 0.000825 | 7.754 | -5.6105 | 0.000825 | 7.754 |  | Radius SMR circle = |  |  | 878.1265 | mm |  | 1.- at each measured position, substract the Zero values from the Raw data of each measured array of three | [16 empty] | 75 |  |  |  | 0.18 |  |  |  | 0.041216 | -6.0763 | -0.029301 |  | 0.17964 |  |  | -0.00035753
| p39 | -870.4406 | 0.89616 | -13.9245 |  | -11.3583 | -0.019095 | 5.1589 | -5.6081 | -0.019095 | 7.7644 |  |  | angle= |  | 0.17 | Deg |  | 2.- remove the tilt induced displacement in X, Y and Z, this delivers the measured commanded displacement values and errors | [16 empty] | 100 |  |  |  | 0.24 |  |  |  | 0.027002 | -8.1081 | -0.03417 |  | 0.23971 |  |  | -0.00028813
| p40 | 853.4485 | -5.7925 | -26.8171 |  | -5.6428 | -5.7788 | -7.7418 | -5.6428 | -0.028635 | -7.7418 |  | offset displacement = |  |  | 2.6055 | mm |  | 3.- at each position define a plane, measure the angle between the commanded and the Zero planes, this yieds the commanded Tilt | [16 empty] | -25 |  |  |  | -0.06 |  |  |  | 0.024608 | 2.0395 | -0.000622 |  | -0.060297 |  |  | -0.00029697
13 | p41 | -5.8978 | 853.2785 | -24.2962 |  | -5.6212 | -5.7805 | -5.2077 | -5.6212 | -0.030338 | -7.8132 | [7 empty] | 4.- from the displacement errors, the Pivot distance and the SMR circle Radius, the Rotation errors were obtained | [16 empty] | -50 |  |  |  | -0.12 |  |  |  | 0.01513 | 4.077 | -0.012543 |  | -0.12053 |  |  | -0.00053404
| p42 | -864.7221 | -4.852 | -26.8922 |  | -5.6398 | -5.7673 | -7.8088 | -5.6398 | -0.017117 | -7.8088 | [7 empty] | 5. from aprox. Position #20 on, much noise on the 3rd level, due to the wash station air- knife blower going on. | [16 empty] | -75 |  |  |  | -0.18 |  |  |  | 0.019829 | 6.1077 | -0.005701 |  | -0.18057 |  |  | -0.00056856
| p43 | 853.4601 | 5.6819 | -26.8395 |  | -5.6312 | 5.6956 | -7.7641 | -5.6312 | -0.054618 | -7.7641 | [24 empty] | -100 |  |  |  | -0.24 |  |  |  | -3.85e-05 | 8.13 | -0.008541 |  | -0.24036 |  |  | -0.0003576
14 | p44 | -5.8977 | 864.7633 | -29.3875 |  | -5.6212 | 5.7043 | -10.299 | -5.6212 | -0.045898 | -7.6935 | [24 empty] | Ry rotational absolute accuracy test
| p45 | -864.7026 | 6.626 | -26.8639 |  | -5.6203 | 5.7108 | -7.7805 | -5.6203 | -0.039398 | -7.7805 | [29 empty] | 0 |  |  | 6.35e-05 | 0 | -1e-06 |  |  | 1.8773e-06 |  | 1.8773e-06
| 31/03/2022, note: the+/- X and Y positions value commanded was 5.600 instead of 5.660mm | [34 empty] | 25% |  |  |  |  | 0.06 |  |  | 2.0254 | 0.017499 | -0.016198 |  |  | 0.05988 |  | -0.00011962
| p1 | 859.0446 | 0 | -19.05 |  | -0.046708 | 0.013646 | 0.025316 | [27 empty] | 50 |  |  |  |  | 0.12 |  |  | 4.0543 | -0.009728 | -0.015929 |  |  | 0.11986 |  | -0.00013625
zero | p2 | -0.3189 | 859.0446 | -19.05 |  | -0.042358 | -0.014446 | 0.038475 | [27 empty] | 75 |  |  |  |  | 0.18 |  |  | 6.0811 | 0.010779 | -0.025934 |  |  | 0.17978 |  | -0.00021592
0.037306 | p3 | -859.0442 | 0.86966 | -19.05 |  | 0.038033 | -0.045603 | 0.033407 | [27 empty] | 100 |  |  |  |  | 0.24 |  |  | 8.1293 | 0.0006175 | -0.033249 |  |  | 0.24034 |  | 0.00033769
| p1 | 853.4146 | 5.7196 | -26.7939 |  | -5.63 | 5.7196 | -7.7439 | -5.63 | -0.030552 | -7.7439 | [24 empty] | -25 |  |  |  |  | -0.06 |  |  | -2.0367 | 0.018046 | -0.010976 |  |  | -0.060215 |  | -0.00021475
14 | p2 | -5.9263 | 864.7616 | -29.3356 |  | -5.6074 | 5.717 | -10.2856 | -5.6074 | -0.033203 | -7.6802 | [24 empty] | -50 |  |  |  |  | -0.12 |  |  | -4.0582 | 0.026688 | 0.011606 |  |  | -0.11998 |  | 2.3465e-05
| p3 | -864.6727 | 6.6114 | -26.8154 |  | -5.6285 | 5.7418 | -7.7653 | -5.6285 | -0.0083852 | -7.7653 | [24 empty] | -75 |  |  |  |  | -0.18 |  |  | -6.0928 | 0.022048 | -0.015558 |  |  | -0.18013 |  | -0.00012997
| p4 | 859.1608 | -0.062168 | -29.3383 |  | 0.11624 | -0.062168 | -10.2883 | -5.6339 | -0.062168 | -7.6828 | [24 empty] | -100 |  |  |  |  | -0.24 |  |  | -8.1132 | 0.013048 | -0.020744 |  |  | -0.23986 |  | 0.000138
15 | p5 | -0.17189 | 859.0177 | -26.798 |  | 0.14701 | -0.02694 | -7.748 | -5.6032 | -0.02694 | -7.748
| p6 | -858.9191 | 0.8419 | -24.3024 |  | 0.1251 | -0.027756 | -5.2524 | -5.6251 | -0.027756 | -7.8579
| p7 | 847.6314 | 0.023668 | -24.2484 |  | -11.4132 | 0.023668 | -5.1984 | -5.663 | 0.023668 | -7.8038 | [23 empty] | IS THIS 3.3.6? | LTS-206, 3.3.4 Maximum Z range (added points + and -8.7mm) | [15 empty] | offset
16 | p8 | -11.7259 | 859.0424 | -26.7848 |  | -11.407 | -0.002235 | -7.7347 | -5.6568 | -0.002235 | -7.7347 | [24 empty] | 3.3.4 |  | Z range in mm | [5 empty] | Raw data with LT in mm |  |  |  | Reduced data mm |  |  | error | removed
| p9 | -870.4478 | 0.84864 | -29.3744 |  | -11.4035 | -0.021012 | -10.3244 | -5.6534 | -0.021012 | -7.7189 | [24 empty] | Max range | X | Y | Z |  |  |  |  | X | Y | Z |  | X | Y | Z | mm | error mm
| p10 | 859.0061 | -0.16748 | -11.2944 |  | -0.038489 | -0.16748 | 7.7556 | -0.038489 | 5.5827 | 7.7556 | [24 empty] | zero | 0 | 0 | 0 |  |  |  |  | -859.0738 | 0.9143 | -19.0653 |  | 0.0002 | 0.0003 | 0 | 0 | 0.019
17 | p11 | -0.34397 | 858.8645 | -8.7633 |  | -0.025072 | -0.18015 | 10.2867 | -0.025072 | 5.57 | 7.6812 | [24 empty] | 25% | 0 | 0 | 1.9 |  |  |  |  | -859.0472 | 0.95659 | -17.2016 |  | 0.026804 | 0.042587 | 1.8637 | -0.036282 | -0.017282
| p12 | -859.08 | 0.71939 | -11.3774 |  | -0.035721 | -0.15027 | 7.6726 | -0.035721 | 5.5999 | 7.6726 | [24 empty] | 50 | 0 | 0 | 3.8 |  |  |  |  | -859.0671 | 0.94424 | -15.2602 |  | 0.006929 | 0.030242 | 3.8051 | 0.005115 | 0.024115
| p13 | 858.9972 | 11.317 | -11.3305 |  | -0.047419 | 11.317 | 7.7195 | -0.047419 | 5.5668 | 7.7195 | [24 empty] | 75 | 0 | 0 | 5.7 |  |  |  |  | -859.033 | 0.96994 | -13.3849 |  | 0.040975 | 0.055939 | 5.6804 | -0.01961 | -0.00061
18 | p14 | -0.36945 | 870.3424 | -13.8812 |  | -0.05055 | 11.2978 | 5.1688 | -0.05055 | 5.5476 | 7.7743 | [24 empty] | 100 | 0 | 0 | 7.6 |  |  |  |  | -859.0501 | 0.95746 | -11.4736 |  | 0.023852 | 0.043462 | 7.5917 | -0.00826 | 0.01074
| p15 | -859.0899 | 12.1881 | -11.3404 |  | -0.045686 | 11.3184 | 7.7096 | -0.045686 | 5.5683 | 7.7096 | [24 empty] | max+ | 0 | 0 | 8.7 |  |  |  |  | -859.0627 | 0.94182 | -10.3637 |  | 0.011344 | 0.027818 | 8.7016 | 0.00161 | 0.02061
[36 empty] | -25 | 0 | 0 | -1.9 |  |  |  |  | -859.0486 | 0.94575 | -21.0097 |  | 0.025355 | 0.031752 | -1.9444 | -0.044371 | -0.025371
19 | [8 empty] | fault and didn't re-measure | [26 empty] | -50 | 0 | 0 | -3.8 |  |  |  |  | -859.0724 | 0.91735 | -22.8661 |  | 0.001627 | 0.003354 | -3.8008 | -0.000768 | 0.018232
[36 empty] | -75 | 0 | 0 | -5.7 |  |  |  |  | -859.0641 | 0.91377 | -24.8138 |  | 0.009862 | -0.000235 | -5.7485 | -0.048456 | -0.029456
| p19 | 853.2128 | 5.5768 | -8.7725 |  | -5.8318 | 5.5768 | 10.2775 | -0.081632 | 5.5768 | 7.6721 | [24 empty] | -100 | 0 | 0 | -7.6 |  |  |  |  | -859.0584 | 0.90435 | -26.6982 |  | 0.015595 | -0.009649 | -7.6329 | -0.032889 | -0.013889
20 | p20 | -6.1334 | 864.6065 | -11.322 |  | -5.8145 | 5.5619 | 7.728 | -0.064308 | 5.5619 | 7.728 | [24 empty] | max- | 0 | 0 | -8.7 |  |  |  |  | -859.065 | 0.89077 | -27.7886 |  | 0.009046 | -0.023235 | -8.7233 | -0.023344 | -0.004344
| p21 | -864.8562 | 6.4515 | -13.8949 |  | -5.812 | 5.5819 | 5.1551 | -0.061804 | 5.5819 | 7.7606 | [25 empty] | 0 | 0 | 0
| p22 | 858.9854 | -0.21527 | -26.7972 |  | -0.059193 | -0.21527 | -7.7472 | -0.059193 | 5.5349 | -7.7472 | [38 empty] | average= | -0.018841 | 0.00015864
21 | p23 | -0.34534 | 858.8417 | -24.2515 |  | -0.026438 | -0.20292 | -5.2015 | -0.026438 | 5.5473 | -7.8069 | [39 empty] | stdev = | 0.019588
| p24 | -859.1022 | 0.68167 | -26.8606 |  | -0.057949 | -0.18798 | -7.8106 | -0.057949 | 5.5622 | -7.8106 | [24 empty] | 30/03/2022 Radial X and Y Translational Range (MOOG 3.3.4), COR= -2758400um
| p25 | 858.9685 | -0.040194 | -19.0648 |  | -0.076114 | -0.040194 | -0.014783 | [28 empty] | Commanded moves to position, from M2 hexapod EUI | [5 empty] | Commanded | Raw data with LT in mm |  |  |  | Reduced data mm |  |  | error | error | error XY
zero | p26 | -0.37876 | 859.0052 | -19.0688 |  | -0.059855 | -0.039457 | -0.018785 | [27 empty] | Move# | X | Y | Z | Rx | Ry | Rz | XY | X | Y | Z |  | X | Y | Z | X | Y | mm
0.021279 | p27 | -859.1109 | 0.82838 | -19.0763 |  | -0.066653 | -0.04128 | -0.026331 | [27 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 |  | 859.0943 | 0.006408 | -19.0566
| p28 | 858.9966 | 11.2704 | -26.8114 |  | 0.028092 | 11.3106 | -7.7467 | 0.028092 | 5.5604 | -7.7467 | [24 empty] | 1 | 7.6 | 0 | 0 | 0 | 0 | 0 | 7.6 | 866.6708 | 0.004087 | -19.0405 |  | 7.5765 | -0.002321 | 0.016095 | -0.023518 | -0.002321 | 0.023632
22 | p29 | -0.33064 | 870.3241 | -29.3542 |  | 0.048121 | 11.319 | -10.2854 | 0.048121 | 5.5688 | -7.6799 | [24 empty] | 2 | 5.374 | 5.374 | 0 | 0 | 0 | 0 | 7.6 | 864.4652 | 5.3596 | -19.0583 |  | 5.3709 | 5.3532 | -0.001718 | -0.003056 | -0.02083 | 0.021053
| p30 | -859.0909 | 12.1944 | -26.8341 |  | 0.020019 | 11.366 | -7.7578 | 0.020019 | 5.6159 | -7.7578 | [24 empty] | 3 | 5.374 | -5.374 | 0 | 0 | 0 | 0 | 7.6 | 864.4601 | -5.3475 | -19.0561 |  | 5.3658 | -5.3539 | 0.000519 | -0.008231 | 0.020113 | 0.021732
| p34 | 864.7152 | 5.4533 | -29.3843 |  | 5.7467 | 5.4935 | -10.3195 | -0.0034632 | 5.4935 | -7.7141 | [24 empty] | 4 | 0 | 7.6 | 0 | 0 | 0 | 0 | 7.6 | 859.1132 | 7.5965 | -19.0378 |  | 0.018893 | 7.5901 | 0.018802 | 0.018893 | -0.009946 | 0.021351
23 | p35 | 5.3774 | 864.5295 | -26.841 |  | 5.7561 | 5.5244 | -7.7722 | 0.0059348 | 5.5244 | -7.7722 | [24 empty] | 5 | -5.374 | 5.374 | 0 | 0 | 0 | 0 | 7.6 | 853.7167 | 5.3766 | -19.0506 |  | -5.3776 | 5.3702 | 0.006003 | -0.003554 | -0.003799 | 0.0052022
| p36 | -853.371 | 6.3832 | -24.3271 |  | 5.7399 | 5.5548 | -5.2508 | -0.010249 | 5.5548 | -7.8562 | [24 empty] | 6 | -7.6 | 0 | 0 | 0 | 0 | 0 | 7.6 | 851.5178 | 0.027719 | -19.0478 |  | -7.5765 | 0.021311 | 0.00878 | 0.023534 | 0.021311 | 0.031749
| p37 | 853.2331 | 5.5 | -24.2571 |  | -5.7354 | 5.5402 | -5.1923 | 0.014793 | 5.5402 | -7.7978 | [24 empty] | 7 | -5.374 | -5.374 | 0 | 0 | 0 | 0 | 7.6 | 853.7241 | -5.3475 | -19.0575 |  | -5.3702 | -5.3539 | -0.000927 | 0.003839 | 0.020113 | 0.020476
24 | p38 | -6.1148 | 864.5519 | -26.8239 |  | -5.7361 | 5.5467 | -7.7551 | 0.014085 | 5.5467 | -7.7551 | [24 empty] | 8 | 0 | -7.6 | 0 | 0 | 0 | 0 | 7.6 | 859.0951 | -7.5789 | -19.0468 |  | 0.000841 | -7.5853 | 0.009793 | 0.000841 | 0.014687 | 0.014711
| p39 | -864.8682 | 6.3839 | -29.4198 |  | -5.7573 | 5.5555 | -10.3435 | -0.0071448 | 5.5555 | -7.738 | [38 empty] | average= |  |  | 0.019988
| p40 | 864.7492 | -5.644 | -13.8663 |  | 5.7807 | -5.6038 | 5.1985 | 5.7807 | 0.14638 | 5.1985 |  |  | 859.0651 | 11.328 | -26.8438 | [33 empty] | stdev = |  |  | 0.0078794
25 | p41 | 5.4202 | 853.4218 | -11.3452 |  | 5.7989 | -5.5834 | 7.7236 | 5.7989 | 0.1668 | 5.1181 | check |  | -0.27287 | 870.4243 | -29.3902 | [19 empty] | 30/03/2022 Rx and Ry rotational Range (MOOG 3.3.8)
| p42 | -853.343 | -4.7536 | -8.887 |  | 5.7679 | -5.582 | 10.1893 | 5.7679 | 0.16817 | 10.1893 |  |  | -859.1074 | 12.3054 | -26.8627 | [20 empty] | Commanded moves to position, from M2 hexapod EUI X: COR= -2758400um; Raw data with LT in mm X: Raw data with LT in mm |  | Reduced data Deg |  |  | error | error | error RxRy
| p43 | 858.9959 | 0.114 | -11.3054 |  | 0.027422 | 0.1542 | 7.7594 | 0.027422 | -5.596 | 7.7594 | [24 empty] | Move# | X | Y | Z | Rx | Ry | Rz |  | X | Y | Z |  | Rx | Ry |  | Rx | Ry | Deg
26 | p44 | -0.34471 | 859.1782 | -13.8709 |  | 0.034043 | 0.17309 | 5.1979 | 0.034043 | -5.5771 | 7.8033 | [24 empty] | pos. zero | 0 | 0 | 0 | 0 | 0 | 0 |  | 1.9e-05 | 2e-06 | -8e-06
| p45 | -859.1061 | 1.0136 | -11.3969 |  | 0.004768 | 0.18523 | 7.6794 | 0.004768 | -5.5649 | 7.6794 | [24 empty] | 1 | 0 | 0 | 0 | 0.24 | 0 | 0 |  | 0.028088 | -8.1071 | -0.007758 |  | 0.23968 | 0.0008304 |  | -0.00032058 | 0.0008304 | 0.00089013
| p46 | 864.743 | -5.6425 | -13.864 |  | 5.7745 | -5.6023 | 5.2008 | 0.024337 | -5.6023 | 7.8062 | [24 empty] | 2 | 0 | 0 | 0 | 0.17 | 0.17 | 0 |  | 5.7878 | -5.7333 | -2.5234 |  | 0.1695 | 0.17111 |  | -0.00049784 | 0.0011127 | 0.001219
27 | p47 | 5.4132 | 853.4212 | -11.3457 |  | 5.792 | -5.5839 | 7.7231 | 0.041776 | -5.5839 | 7.7231 | [24 empty] | 3 | 0 | 0 | 0 | 0 | 0.24 | 0 |  | 8.1235 | 0.025463 | -3.591 |  | -0.0007528 | 0.24017 |  | -0.0007528 | 0.00016697 | 0.00077109
| p48 | -853.3501 | -4.738 | -8.8746 |  | 5.7608 | -5.5663 | 10.2018 | 0.01058 | -5.5663 | 7.5963 | [24 empty] | 4 | 0 | 0 | 0 | -0.17 | 0.17 | 0 |  | 5.7476 | 5.7708 | -2.5382 |  | -0.17061 | 0.16992 |  | -0.00060967 | -7.6075e-05 | 0.0006144
| p49 | 853.2331 | -5.6692 | -8.7949 |  | -5.7354 | -5.629 | 10.2699 | 0.014766 | -5.629 | 7.6645 | [24 empty] | 5 | 0 | 0 | 0 | -0.24 | 0 | 0 |  | -0.009537 | 8.1234 | -0.020914 |  | -0.24016 | -0.00028195 |  | -0.00016357 | -0.00028195 | 0.00032596
28 | p50 | -6.1166 | 853.4107 | -11.3599 |  | -5.7379 | -5.5945 | 7.7089 | 0.012283 | -5.5945 | 7.7089 | [24 empty] | 6 | 0 | 0 | 0 | -0.17 | -0.17 | 0 |  | -5.7504 | 5.7752 | 2.5421 |  | -0.17074 | -0.17001 |  | -0.00073836 | -6.4982e-06 | 0.00073839
| p51 | -864.8633 | -4.7896 | -13.9884 |  | -5.7524 | -5.618 | 5.0879 | -0.0021928 | -5.618 | 7.6934 | [24 empty] | 7 | 0 | 0 | 0 | 0 | -0.24 | 0 |  | -8.1178 | 0.049969 | 3.6131 |  | -0.0014773 | -0.24 |  | -0.0014773 | 1.6379e-06 | 0.0014773
| p52 | 858.9672 | -11.4408 | -26.9209 |  | 0.005466 | -11.3244 | -7.8084 | 0.005466 | -5.5742 | -7.8084 | check |  | 859.0668 | 0.077965 | -26.8342 | [19 empty] | 8 | 0 | 0 | 0 | 0.17 | -0.17 | 0 |  | -5.7532 | -5.745 | 2.5543 |  | 0.16985 | -0.17009 |  | -0.00015318 | -8.9515e-05 | 0.00017742
29 | p53 | -0.35537 | 847.6221 | -24.3061 |  | 0.065925 | -11.3377 | -5.199 | 0.065925 | -5.5876 | -7.8044 |  |  | -0.30415 | 859.1573 | -29.409 | [33 empty] | average= |  |  | 0.00077671
| p54 | -859.1046 | -10.5408 | -26.9209 |  | 0.028078 | -11.3229 | -7.8092 | 0.028078 | -5.5728 | -7.8092 |  |  | -859.1256 | 0.99102 | -26.9007 | [33 empty] | stdev = |  |  | 0.0004292
| p55 | 858.9672 | 0.074413 | -26.8287 |  | 0.005466 | 0.19076 | -7.7161 | 0.005466 | -5.5594 | -7.7161
30 | p56 | -0.38731 | 859.1389 | -29.3927 |  | 0.03398 | 0.17908 | -10.2856 | 0.03398 | -5.5711 | -7.6801
| p57 | -859.1342 | 0.94369 | -26.9008 |  | -0.001548 | 0.16151 | -7.7891 | -0.001548 | -5.5887 | -7.7891
| p58 | 864.6757 | -5.7042 | -29.3993 |  | 5.714 | -5.5879 | -10.2868 | -0.036219 | -5.5879 | -7.6813
31 | p59 | 5.3745 | 853.3577 | -26.8777 |  | 5.7957 | -5.6022 | -7.7706 | 0.045571 | -5.6022 | -7.7706
| p60 | -853.4036 | -4.7964 | -24.3627 |  | 5.729 | -5.5786 | -5.251 | -0.021129 | -5.5786 | -7.8565
| p61 | 853.2144 | -5.6585 | -24.2643 |  | -5.7473 | -5.5422 | -5.1517 | 0.0028292 | -5.5422 | -7.7572
32 | p62 | -6.1227 | 853.3679 | -26.8376 |  | -5.7014 | -5.5919 | -7.7306 | 0.04874 | -5.5919 | -7.7306
| p63 | -864.8951 | -4.7907 | -29.4176 |  | -5.7625 | -5.5729 | -10.3058 | -0.0123 | -5.5729 | -7.7003
| p64 | 858.9618 | -0.11635 | -19.1125 |  | -0.006711 | -0.076151 | -0.047753
zero | p65 | -0.42129 | 858.9598 | -19.1071 |  | -0.042536 | -0.04531 | -0.038278
0.019058 | p66 | -859.1326 | 0.78218 | -19.1118 |  | -0.021742 | -0.046196 | -0.035413
Error
of zeros
stdev
0.019136 | mm

## ResZ 20210614
Index | [16 empty] | measured February 2022 with comcam mounted | [8 empty] | Repeated Feb. Measurements |  |  | measured
meas | COR= -2758400um (pivot point on vertex L1S1) | [11 empty] | expected |  |  | Raw measurement |  |  |  | meas | measured | resol |  |  | calib | mitutoyo linear gauges values |  |  | offset | error
step | Commanded | EUI | Mitutoyo Offset | Mitutoyo Raw | Mitutoyo Value | [6 empty] | error | gauge |  |  | Mitutoyo linear gauge |  |  |  | step | offset with | error |  |  | offset | X | Y | Z | Z | %
# | Offset Z | Z value | mm | mm | mm | 1 | 2 | 3 | 4 | 5 | 6 | % | um | x | y | z |  |  |  |  | gauge | Zm-Zset |  |  | Z | mm | mm | mm | um | of move
| 0 | 0 | -4.655 | -4.655 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 |  |  | 0.013 | 0.02 |  | 0.014 |  | 0 |  |  | 0 | 0.012 | -0.005 | 0.019 | 0.001
1 | 5 | 5 |  | -4.65 | 0.005 | [6 empty] | 2.1316e-12 | 5 |  |  | 0.02 | 0.026 |  |  | 1 | 0.006 | 0.001 |  | 50 | 0.062 | -0.005 | 0.021 | 0.007 | 6 | 20 | bar: 1 to 6
2 | 5 | 11 |  | -4.646 | 0.009 | [6 empty] | 40 | 10 |  |  | 0.025 | 0.031 |  |  | 2 | 0.011 | 0 |  | 0 | 0.012 | -0.005 | 0.02 | 0.012 | 5 | 0 | bar: 1 to 6
3 | -5 | 5 |  | -4.65 | 0.005 | [6 empty] | 40 |  |  |  | 0.019 | 0.025 |  |  | 3 | 0.005 | 0 |  |  |  | -0.006 | 0.019 | 0.007 | -5 | 0 | bar: 1 to 3; 5 to 6
4 | 5 | 11 |  | -4.646 | 0.009 | [6 empty] | 40 | 5 |  |  | 0.025 | 0.032 |  |  | 4 | 0.012 | 0.001 |  |  |  | -0.006 | 0.021 | 0.011 | 4 | -20 | bar: 1 to 6
5 | -5 | 5 |  | -4.65 | 0.005 | [6 empty] | 40 | 0 |  |  | 0.019 | 0.026 |  |  | 5 | 0.006 | 0.001 |  |  |  | -0.006 | 0.019 | 0.006 | -5 | -1.7764e-14 | bar: 1 to 6
|  | 5 | -4.65 | -4.65 | 0 | [11 empty] | 0.026 |  |  |  | 0.006 | 0.001 |  |  |  | -0.006 | 0.02 | 0.002
1 | 2 | 8 |  | -4.648 | 0.002 | [6 empty] | 50 | 2 |  |  |  | 0.028 |  |  | 1 | 0.008 | 0 | 2 |  |  | -0.005 | 0.021 | 0.004 | 2 | 0 | bar: 1 to 6
2 | 2 | 10 |  | -4.646 | 0.004 | [6 empty] | 1.1013e-11 | 4 |  |  |  | 0.03 |  |  | 2 | 0.01 | 0 | 2 |  |  | -0.005 | 0.022 | 0.006 | 2 | 0 | bar: 1 to 6
3 | -2 | 8 |  | -4.648 | 0.002 | [6 empty] | 1.1013e-11 | 2 |  |  |  | 0.029 |  |  | 3 | 0.009 | 0.001 | -1 |  |  | -0.005 | 0.021 | 0.005 | -1 | -50 | bar: 1 to 6
4 | 2 | 10 |  | -4.646 | 0.004 | [6 empty] | 1.1013e-11 | 4 |  |  |  | 0.03 |  |  | 4 | 0.01 | 0 | 1 |  |  | -0.005 | 0.021 | 0.006 | 1 | -50 | bar: 1 to 6
5 | -2 | 8 |  | -4.648 | 0.002 | [6 empty] | 1.1013e-11 | 2 |  |  |  | 0.029 |  |  | 5 | 0.009 | 0.001 | -1 |  |  | -0.005 | 0.02 | 0.005 | -1 | -50 | bar: 1 to 6
|  | 8 | -4.648 | -4.648 | 0 | [11 empty] | 0.029 | 0.031 | 0.036 |  | 0 | [8 empty] | 30
1 | 1 | 9 |  | -4.647 | 0.001 | [6 empty] | 5.5422e-11 | 1 |  |  |  | 0.029 | 0.032 | 0.036 | 1 | 0.001 | 0 | 1 |  |  | -0.005 | 0.02 | 0.005 | bar: 1 to 6
2 | 1 | 10 |  | -4.646 | 0.002 | [6 empty] | 3.3396e-11 | 2 |  |  |  | 0.029 | 0.033 | 0.037 | 2 | 0.002 | 0.001 | 1 |  |  | -0.005 | 0.02 | 0.005 | these didm't work | bar: 1 to 6
3 | -1 | 9 |  | -4.647 | 0.001 | [6 empty] | 3.3396e-11 | 1 |  |  |  | 0.029 | 0.033 | 0.037 | 3 | 0.002 | 0.001 | 0 |  |  | -0.005 | 0.02 | 0.005 | bar: 1 to 6
4 | 1 | 10 |  | -4.646 | 0.002 | [6 empty] | 3.3396e-11 | 2 |  |  |  | 0.029 | 0.033 | 0.037 | 4 | 0.002 | 0.001 | 0 |  |  | -0.005 | 0.02 | 0.005 | bar: 1 to 6
5 | -1 | 9 |  | -4.647 | 0.001 | [6 empty] | 3.3396e-11 | 1 |  |  |  | 0.29 | 0.033 | 0.037 | 5 | 0.002 | 0.001 | 0 |  |  | -0.006 | 0.019 | 0.005 | bar: 1 to 6
[6 empty] | 6: av.error%= | 14 |  |  |  |  | 0.031 | [8 empty] | -0.006 | 0.017 | 0.005 | offsets in Z
[6 empty] | 6: stdev%= | 20.6328 | [13 empty] | -0.006 | 0.016 | 0.005 | -1
| repeated in Feb2022 | [24 empty] | -0.006 | 0.015 | 0.005 | -1
| 1um offset | [24 empty] | -0.006 | 0.014 | 0.004 | -1
set | zero in abs.mode | [24 empty] | -0.006 | 0.022 | 0.002 | abs. mode set zero
1 | 1 | [24 empty] | -0.006 | 0.022 | 0.003 | 1 | 0
2 | 1 | [24 empty] | -0.006 | 0.022 | 0.004 | 1 | 0
3 | -1 | [24 empty] | -0.006 | 0.022 | 0.004 | 0 | -100
4 | 1 | [24 empty] | -0.006 | 0.022 | 0.005 | 1 | 0
5 | -1 | [24 empty] | -0.006 | 0.021 | 0.005 | 0 | -100
40

## ResY 20210615
Index | COR= -2758400um (pivot point on L1S1) | [15 empty] | measured w/Mitutoyos Feb 2022 |  |  |  |  | Offset | error
Meas. |  |  | display | Mitutoyo linear gauge Y | [8 empty] | expected |  | Y value | [6 empty] | measured | %
step | Commanded | EUI | EUI | Offset | Y Raw | Y Value | EUI Hexapod Strut position e.u (um) | [5 empty] | gauge | error | EUI | [7 empty] | of move
| Offset Yum | Y um | offset | mm | mm | mm | 1 | 2 | 3 | 4 | 5 | 6 | um | % |  | X | Y | Z |  |  |  | Y (um)
| 0 | -200 |  | -0.204 |  |  | -58 | -57 | 114 | -57 | -58 | 115 |  |  |  | -0.005 | 0.014 | -0.006
1 | 5 | -195 | 5 |  | -0.197 | 0.007 | -56 | -56 | 111 | -56 | -56 | 111 | 5 | 40 | 4 | 0.003 | 0.02 | -0.004 |  |  |  | 6 | 20
2 | 5 | -190 | 10 |  | -0.191 | 0.013 | -54 | -54 | 108 | -54 | -55 | 108 | 10 | 60 | 10 | -0.003 | 0.025 | -0.005 |  |  |  | 5 | 1.7764e-14
3 | -5 | -195 | 5 |  | -0.198 | 0.006 | -56 | -55 | 111 | -56 | -56 | 111 | 5 | 20 | 4 | -0.005 | 0.019 | -0.005 |  |  |  | -6 | 20
4 | 5 | -190 | 10 |  | -0.191 | 0.013 | -54 | -54 | 108 | -54 | -55 | 108 | 10 | 60 | 9 | -0.005 | 0.024 | -0.003 |  |  |  | 5 | 1.7764e-14
5 | -5 | -195 | 5 |  | -0.198 | 0.006 | -56 | -55 | 111 | -56 | -56 | 111 | 5 | 20 | 3 | -0.005 | 0.017 | -0.004 |  |  |  | -7 | 40
|  |  |  | -0.198 | [10 empty] | 3 | [7 empty] | 16
1 | 2 | -193 | 2 |  | -0.196 | 0.002 | -55 | -55 | 109 | -55 | -56 | 110 | 2 | 8.8818e-14 | 4 | -0.005 | 0.019 | -0.004 |  |  |  | 2 | -8.8818e-14
2 | 2 | -190 | 5 |  | -0.193 | 0.005 | -55 | -54 | 108 | -55 | -55 | 109 | 4 | 50 | 5 | -0.005 | 0.021 | -0.004 |  |  |  | 2 | 8.8818e-14
3 | -2 | -193 | 2 |  | -0.195 | 0.003 | -56 | -55 | 109 | -55 | -56 | 110 | 2 | 50 | 2 | -0.006 | 0.019 | -0.004 | comment: X value in EUI is -2um |  |  | -2 | 8.8818e-14
4 | 2 | -191 | 4 |  | -0.194 | 0.004 | -55 | -54 | 108 | -55 | -55 | 109 | 4 | 1.7764e-13 | 4 | -0.006 | 0.02 | -0.003 |  | x=-3 |  | 1 | -50
5 | -2 | -193 | 2 |  | -0.195 | 0.003 | -55 | -55 | 109 | -55 | -56 | 110 | 2 | 50 | 1 | -0.007 | 0.018 | -0.003 |  | x=-2 |  | -2 | 8.8818e-14
| 0 | [11 empty] | av.error%= | 35 | -2 |  | 0.015 |  |  | x=-2 | Z=-3um
[13 empty] | stdev%= | 23.214
-0.195
| 1 | -192 |  |  | -0.195 | 0 | -55 | -55 | 108 | -55 | -56 | 110
| 1 | -191 |  |  | -0.194 | 0.001 | -55 | -55 | 107 | -55 | -56 | 109
| -1 | -192 |  |  | -0.194 | 0.001 | -55 | -55 | 108 | -55 | -56 | 110
| 1 | -191 |  |  | -0.194 | 0.001 | -55 | -55 | 107 | -55 | -56 | 109
| -1 | -192 |  |  | -0.194 | 0.001 | -56 | -55 | 108 | -55 | -56 | 110

## ResX COR -2758400um
|  |  | Gauge X(Rz) | [20 empty] | measured w/Mitutoyos Feb.2022
[28 empty] | measurements |  |  |  | error
Meas. | COR= -2758400um (Pivot point at L1S1) | [18 empty] | expected |  | Xvalue | [5 empty] | 2022-02-01 |  |  | error | %
offset | Commanded | EUI |  | Mitutoyo linear gauges |  |  |  | EUI Hexapod Strut position e.u. (um) | [6 empty] | hexapod EUI display (um) and (Deg) |  |  |  |  | gauge | error | EUI | [8 empty] | um | of move
Index # | Offset X | X value | X Raw | X Value | Y | Zx+ | Zy- | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | um | % |  | X | Y | Z | Delta T | 0.04C | X (um) | Y (um) | Z (um) | X | X
|  | -50 | 0.202 |  | 0.073 | 3.122 | 0.052 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -0.009 | 0.013 | -0.005
1 | 5 | -45 | 0.209 | 0.007 | 0.073 | 3.122 | 0.052 | -3 | 3 | 0 | -3 | 3 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 5 | 40 | 6 | -0.004 | 0.013 | -0.004 |  |  | 5 | 0 | 1 | 0 | 0
2 | 5 | -39 | 0.214 | 0.012 | 0.073 | 3.122 | 0.052 | -6 | 6 | 0 | -6 | 5 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 10 | 40 | 11 | 0.001 | 0.013 | -0.004 |  |  | 5 | 0 | 0 | 0 | 0
3 | -5 | -45 | 0.21 | 0.008 | 0.073 | 3.122 | 0.052 | -3 | 3 | 0 | -3 | 3 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 5 | 60 | 6 | -0.003 | 0.013 | -0.004 |  |  | -4 | 0 | 0 | 1 | -20
4 | 5 | -39 | 0.214 | 0.012 | 0.073 | 3.122 | 0.052 | -6 | 6 | 0 | -6 | 5 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 10 | 40 | 11 | 0.001 | 0.013 | -0.004 |  |  | 4 | 0 | 0 | -1 | -20
5 | -5 | -45 | 0.21 | 0.008 | 0.073 | 3.122 | 0.052 | -4 | 3 | 0 | -3 | 2 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 5 | 60 | 6 | -0.003 | 0.013 | -0.004 |  |  | -4 | 0 | 0 | 1 | -20
|  | -40 | 0.213 |  | 0.085 | 3.116 | 0.056 | 34 | -33 | 0 | 33 | -33 | 0 | -67 | 0 | 0 | 0 | 0 | 0 |  |  | 6 | [9 empty] | -12
1 | 2 | -37 | 0.216 | 0.003 | 0.085 | 3.116 | 0.056 | 32 | -32 | 0 | 32 | -32 | 0 | -65 | 0 | 0 | 0 | 0 | 0 | 2 | 20 | 8 | -0.003 | 0.013 | -0.004 |  |  | 0 | 0 | 0 | -2 | -100
2 | 2 | -35 | 0.218 | 0.005 | 0.085 | 3.116 | 0.056 | 31 | -30 | 0 | 31 | -31 | 0 | -62 | 0 | 0 | 0 | 0 | 0 | 4 | 20 | 11 | -0.001 | 0.013 | -0.004 |  |  | 2 | 0 | 0 | 0 | 0
3 | -2 | -37 | 0.218 | 0.005 | 0.085 | 3.116 | 0.056 | 32 | -31 | 0 | 32 | -32 | 0 | -64 | -1 | 0 | 0 | 0 | 0 | av.error%= | 40 | 9 | -0.001 | 0.012 | -0.004 |  |  | 0 | -1 | 0 | 2 | -100
4 | 2 | -35 | 0.284 | 0.071 | 0.085 | 3.116 | 0.056 | [12 empty] | stdev%= | 16.3299 | 11 | -0.001 | 0.012 | -0.004 |  |  | 0 | 0 | 0 | -2 | -100
5 | -2 | -37 | 0.117 | -0.096 | [17 empty] | 9 | -0.001 | 0.012 | -0.004 | screen shot of status here |  | 0 | 0 | 0 | 2 | -100
|  | -47 | 0.107 | [22 empty] | (before going back to zero)
| 1 | -46 | 0.109 | 0.002
| 1 | -45 | 0.109 | 0.002
| -1 | -46 | 0.109 | 0.002
| 1 | -45 | 0.11 | 0.003
| -1 | -46 | 0.109 | 0.002

## ResRx COR -2758400um
Meas. | [24 empty] | measured w/mitutoyos Feb 2022
Offset | COR= -2758400um (Pivot point at L1S1) | [19 empty] | expected |  |  | Rx |  |  | res offset= | 8.19e-05 | Deg
Index |  | Commanded | EUI | (8.19E-5 deg<=> 2.7702um of Y Gauge) |  |  |  |  | EUI Hexapod Strut position |  |  |  | e.u. (um) |  |  |  | hexapod EUI display |  |  |  | gauge | error |  | value |  | Rx+ => | Y shift = | -0.0027702 | mm
# |  | Offset Rx | Rx value | Y raw | Y value | G. X (Rz) | Gauge Zx | gaugeZy- | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | um | % |  | EUI |  | Rx+ => | Z shift = | -0.0012436 | mm
| 5x | (5x 8.19E-5) | 0 | 0.069 |  | 0.167 | 3.121 | 0.052 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | Fault 10:35. Drive Fault, Simulik Fault
|  | 0.0004095 | 0.0005 | 0.055 | -0.014 | 0.162 | 3.118 | 0.046 | -9 | -9 | 10 | -2 | -1 | 10 | 0 | 1 | 0 | 0.0004 | 0 | 0 | 13.851 | 2.98
|  | 0.0004095 | 0.0009 | 0.042 | -0.027 | 0.162 | 3.118 | 0.04 | -18 | -18 | 20 | -3 | -2 | 20 | -1 | 2 | 0 | 0.0009 | 0 | 0 | 27.702 | 14.04
|  | -0.0004095 | 0.0005 | 0.056 | -0.013 | 0.164 | 3.119 | 0.046 | -9 | -9 | 10 | -2 | -1 | 10 | -1 | 1 | 0 | 0.0004 | 0 | 0 | 13.851 | 17.02
|  | 0.0004095 | 0.0009 | 0.043 | -0.026 | 0.163 | 3.117 | 0.04 | -18 | -17 | 20 | -3 | -3 | 20 | -1 | 2 | 0 | 0.0009 | 0 | 0 | 27.702 | 34.04
|  | -0.0004095 | 0.0005 | 0.056 | -0.013 | 0.164 | 3.118 | 0.046 | -9 | -9 | 9 | -2 | -1 | 10 | -1 | 1 | 0 | 0.0004 | 0 | 0 | 13.851 | 17.02
| 2x | (2x 8.19E-5) | 0.0005 | 0.056 |  | 0.164 | 3.118 | 0.046 | -9 | -9 | 9 | -2 | -1 | 10 | -1 | 1 | 0 | 0.0004 | 0 | 0
|  | 0.0001638 | 0.0007 | 0.05 | -0.006 | 0.164 | 3.117 | 0.043 | -13 | -13 | 13 | -2 | -2 | 15 | -2 | 2 | 0 | 0.0006 | 0 | 0 | 5.5404 | 9.192
|  | 0.0001638 | 0.0009 | 0.045 | -0.011 | 0.163 | 3.116 | 0.041 | -16 | -16 | 17 | -3 | -2 | 19 | -2 | 3 | 0 | 0.0008 | 0.0001 | 0 | 11.0808 | 1.616 | [6 empty] | error |  | error
|  | -0.0001638 | 0.0007 | 0.052 | -0.004 | 0.163 | 3.119 | 0.043 | -13 | -12 | 13 | -2 | -2 | 15 | -1 | 2 | 0 | 0.0006 | 0 | 0 | 5.5404 | 30.808 | [5 empty] | Rx from | as % | Rx from | as %
|  | 0.0001638 | 0.0009 | 0.046 | -0.01 | 0.163 | 3.116 | 0.041 | -17 | -16 | 17 | -3 | -2 | 19 | -2 | 3 | 0 | 0.0008 | 0.0001 | 0 | 11.0808 | 21.616 | [5 empty] | meas. Y | of Y move | meas. Z | of z move
|  | -0.0001638 | 0.0007 | 0.052 | -0.004 | 0.163 | 3.119 | 0.043 | -13 | -12 | 13 | -2 | -2 | 15 | -1 | 2 | 0 | 0.0006 | 0 | 0 | 5.5404 | 30.808 |  |  | X | Y | Z | (Deg) | % | (Deg) | %
| 1x | (1x 8.19E-5) | 0.0004 | 0.052 |  | 0.163 | 3.119 | 0.043 | -13 | -12 | 13 | -2 | -2 | 15 | -1 | 2 | 0 | 0.0006 | 0 | 0 |  |  |  | 0.000112 | -0.008 | 0.01 | -0.006
1 | 8.19e-05 | 8.19e-05 | 0.0005 | 0.049 | -0.003 | 0.163 | 3.117 | 0.042 | -15 | -14 | 16 | -2 | -2 | 17 | -2 | 3 | 0 | 0.0007 | 0.0001 | 0 | 2.7702 | 4.596 |  | 0.0002 | -0.008 | 0.006 | -0.008 | 0.00011826 | 44.3923 | 0.00013096 | 59.9041
2 | 8.19e-05 | 8.19e-05 | 0.0006 | 0.046 | -0.006 | 0.163 | 3.116 | 0.04 | -17 | -16 | 18 | -2 | -2 | 20 | -2 | 4 | 0 | 0.0009 | 0.0001 | 0 | 5.5404 | 9.192 |  | 0.0004 | -0.009 | 0.003 | -0.011 | 8.8693e-05 | 8.2942 | 0.00019644 | 139.8562
3 | -8.19e-05 | -8.19e-05 | 0.0005 | 0.049 | -0.003 | 0.163 | 3.117 | 0.041 | -15 | -14 | 16 | -2 | -2 | 18 | -3 | 3 | 0 | 0.0008 | 0.0001 | -0.0001 | 2.7702 | 4.596 |  | 0.0002 | -0.008 | 0.008 | -0.008 | -0.00014782 | 80.4903 | -0.00019644 | 139.8562
4 | 8.19e-05 | 8.19e-05 | 0.0006 | 0.046 | -0.006 | 0.163 | 3.116 | 0.04 | -17 | -16 | 18 | -2 | -2 | 20 | -3 | 4 | 0 | 0.0009 | 0.0001 | -0.0001 | 5.5404 | 9.192 |  | 0.000354 | -0.008 | 0.004 | -0.01 | 0.00011826 | 44.3923 | 0.00013096 | 59.9041
5 | -8.19e-05 | -8.19e-05 | 0.0005 | 0.048 | -0.004 | 0.163 | 3.117 | 0.042 | -15 | -14 | 16 | -2 | -2 | 18 | -3 | 3 | 0 | 0.0008 | 0.0001 | -0.0001 | 2.7702 | 24.596 |  | 0.000251 | -0.008 | 0.008 | -0.007 | -0.00011826 | 44.3923 | -0.00019644 | 139.8562
| repeatx1 | 0 |  | 0.07 |  | 0.164 | 3.12 | 0.052 | 0 | 0 | -1 | 0 | 0 | -1 | 0 | -1 | 0 | 0 | 0 | 0 | [8 empty] | 44.3923 |  | 107.8754
|  | 8.19e-05 | 0.0001 | 0.067 | -0.003 | 0.164 | 3.118 | 0.05 | -2 | -2 | 3 | 0 | -1 | 1 | 0 | 0 | 0 | 0.0001 | 0 | 0 | 2.7702 | 4.596
|  | 8.19e-05 | 0.0002 | 0.064 | -0.006 | 0.164 | 3.116 | 0.048 | -4 | -4 | 4 | 0 | -1 | 4 | -1 | 1 | 0 | 0.0002 | 0 | 0 | 5.5404 | 9.192
|  | -8.19e-05 | 0.0001 | 0.067 | -0.003 | 0.164 | 3.117 | 0.049 | -2 | -2 | 2 | 0 | -1 | 2 | -1 | 0 | 0 | 0.0001 | 0 | 0 | 2.7702 | 4.596
|  | 8.19e-05 | 0.0002 | 0.064 | -0.006 | 0.164 | 3.116 | 0.048 | -4 | -4 | 4 | 0 | -1 | 4 | -1 | 1 | 0 | 0.0002 | 0 | 0 | 5.5404 | 9.192
|  | -8.19e-05 | 0.0001 | 0.067 | -0.003 | 0.164 | 3.117 | 0.049 | -2 | -2 | 2 | 0 | -1 | 2 | -1 | 0 | 0 | 0.0001 | 0 | 0 | 2.7702 | 4.596
| zero off.mode | -8.19e-05 | 0 | 0.07 |  | 0.164 | 3.119 | 0.051 | 0 | 0 | -1 | 0 | 0 | 0 | -1 | -1 | 0 | 0 | 0 | 0
| in abs.mode | Set to 0 | 0 | 0.07 |  | 0.164 | 3.119 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0
[21 empty] | av.error%= | 13.1742
[21 empty] | stdev%= | 10.193
[5 empty] | 0.01385 |  | 0.00277
0.0027702

## ResRy COR -2758400um
offset | [31 empty] | measured w/Mitutoyos Feb.2022 | [10 empty] | res offset | 8.19e-05 | deg
index | COR= -2758400um (Pivot point at L1S1) | [19 empty] | expected |  | start | Ry |  |  | measured |  |  | measured |  | Ry from | error as | RX from | RxRy | error as |  | Ry from |  | improve setup
# |  | Commanded | EUI Ry | (Ry=8.19E-5 deg<=> 0.996845um of Zx+ Gauge) |  |  |  |  | EUI Hexapod Strut position |  |  |  | e.u. (um) |  |  |  | hexapod EUI display |  |  |  | gauge | error | position | value |  |  | x offsets |  |  | Y offsets |  | X meas. | % of | Y meas. |  | % of | XY | XY meas.
|  | Offset Ry | Deg | Zx+ raw | Zx+ Value | Gauge Y | G. X (Rz) | Gauge Zy- | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | um | % |  | EUI | X |  | mm | Y |  | mm | Z | (Deg) | move | (Deg) |  | move |  | (Deg) |  | X | Y | Z
| 5x | (5x 8.19E-5) | 0 | 3.119 | 0 | 0.071 | 0.166 | 0.05 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
1 |  | 0.0004095 | 0.0005 | 3.11 | -0.009 | 0.068 | 0.178 | 0.051 | -7 | 7 | -4 | -11 | 10 | 4 | 0 | -1 | 0 | 0 | 0.0004 | 0 | 4.9842 | 80.3155
2 |  | 0.0004095 | 0.0009 | 3.104 | -0.015 | 0.067 | 0.19 | 0.051 | -14 | 14 | -8 | -22 | 21 | 9 | 0 | -1 | 0 | 0 | 0.0008 | 0 | 9.9685 | 50.3155
3 |  | -0.0004095 | 0.0005 | 3.113 | -0.006 | 0.071 | 0.178 | 0.05 | -7 | 7 | -4 | -11 | 10 | 4 | 0 | -1 | 0 | 0 | 0.0004 | 0 | 4.9842 | 20.3155
4 |  | 0.0004095 | 0.0009 | 3.104 | -0.015 | 0.068 | 0.19 | 0.051 | -14 | 14 | -8 | -22 | 21 | 9 | 0 | -1 | 0 | 0 | 0.0008 | 0 | 9.9685 | 50.3155
5 |  | -0.0004095 | 0.0005 | 3.112 | -0.007 | 0.071 | 0.178 | 0.05 | -7 | 7 | -4 | -11 | 10 | 4 | 0 | -1 | 0 | 0 | 0.0004 | 0 | 4.9842 | 40.3155
| 2x | (2x 8.19E-5) | 0.0005 | 3.112 | 0 | 0.071 | 0.178 | 0.05 | -7 | 7 | -4 | -11 | 10 | 4 | 0 | -1 | 0 | 0 | 0.0004 | 0 |  |  |  |  | -0.006 |  |  | 0.014 |  |  | 0.001
1 |  | 0.0001638 | 0.0007 | 3.109 | -0.003 | 0.071 | 0.183 | 0.05 | -10 | 10 | -6 | -15 | 15 | 5 | 1 | -1 | 0 | 0 | 0.0006 | 0 | 1.9937 | 50.3155 |  |  | -0.001 | 0.005 | 0.005 | 0.012 | -0.002 | -0.002 | 0 | 0.00014782 | -9.7548 | -5.9129e-05 | 0.00015921 | -2.803 | 0.0053852 | 0.00015921
2 |  | 0.0001638 | 0.0009 | 3.106 | -0.006 | 0.069 | 0.188 | 0.051 | -13 | 13 | -7 | -20 | 19 | 7 | 1 | -1 | 0 | 0 | 0.0007 | 0 | 3.9874 | 50.3155 |  |  | 0.009 | 0.015 | 0.01 | 0.01 | -0.004 | -0.002 | -0.001 | 0.00029564 | 80.4903 | -5.9129e-05 | 0.0003015 | 84.0647 | 0.015524 | 0.00029975
3 |  | -0.0001638 | 0.0007 | 3.109 | -0.003 | 0.071 | 0.183 | 0.051 | -10 | 10 | -6 | -15 | 15 | 6 | 1 | -1 | 0 | 0 | 0.0006 | 0 | 1.9937 | 50.3155 |  |  | 0.002 | 0.008 | -0.007 | 0.014 | 0 | 0.004 | 0 | -0.00020695 | 26.3432 | 0.00011826 | 0.00023836 | -45.516 | 0.008 | -0.00022245
4 |  | 0.0001638 | 0.0009 | 3.106 | -0.006 | 0.07 | 0.187 | 0.051 | -13 | 13 | -7 | -20 | 19 | 7 | 1 | -1 | 0 | 0 | 0.0007 | 0 | 3.9874 | 50.3155 |  |  | 0.01 | 0.016 | 0.008 | 0.01 | -0.004 | -0.004 | -0.001 | 0.00023651 | 44.3923 | -0.00011826 | 0.00026443 | 61.4355 | 0.016492 | 0.00025107
5 |  | -0.0001638 | 0.0007 | 3.109 | -0.003 | 0.07 | 0.183 | 0.051 | -10 | 10 | -6 | 16 | 15 | 6 | 1 | -1 | 0 | 0 | 0.0006 | 0 | 1.9937 | 50.3155 |  |  | 0.002 | 0.008 | -0.008 | 0.014 | 0 | 0.004 | 0 | -0.00023651 | 44.3923 | 0.00011826 | 0.00026443 | -61.4355 | 0.008 | -0.00025107
| 1x | (1x 8.19E-5) | 0.0004 | 3.109 | 0 | 0.07 | 0.183 | 0.051 | -10 | 10 | -6 | 16 | 15 | 6 | 1 | -1 | 0 | 0 | 0.0006 | 0 | [12 empty] | 41.0746 |  |  | 7.1492 |  |  | 1.8e-06 |  | -0.005 | 0.014 | -0.002 |  | 0 | -0.005 | 0.014 | 0
1 | 8.19e-05 | 8.19e-05 | 0.0005 | 3.108 | -0.001 | 0.069 | 0.185 | 0.05 | -12 | 11 | -7 | -18 | 17 | 6 | 1 | -1 | 0 | 0 | 0.0007 | 0 | 0.99684 | 0.3155 |  |  | 0.007 | 0.013 | 0.005 | 0.014 | 0 | 0 | 0 | 0.00014782 | 80.4903 | 0 | 0.00014782 | 80.4903 | 0.015652 |  | 0.0002 | -0.003 | -0.004 | 0.014 | -0.002 |  |  | -0.005 | 0.014 | 0
2 | 8.19e-05 | 8.19e-05 | 0.0006 | 3.107 | -0.002 | 0.069 | 0.188 | 0.05 | -13 | 13 | -7 | -20 | 19 | 7 | 2 | -1 | 0 | 0 | 0.0007 | 0 | 1.9937 | 0.3155 |  |  | 0.009 | 0.015 | 0.002 | 0.013 | -0.001 | -0.001 | 0 | 5.9129e-05 | -27.8039 | -2.9564e-05 | 6.6108e-05 | -19.2823 | 0.015811 |  | 0.0004 |  | -0.003 | 0.014 | -0.002 |  |  | -0.004 | 0.014 | 0
3 | -8.19e-05 | -8.19e-05 | 0.0005 | 3.108 | -0.001 | 0.069 | 0.188 | 0.05 | -12 | 11 | -7 | -18 | 17 | 6 | 2 | -1 | 0 | 0 | 0.0006 | 0 | 0.99684 | 0.3155 |  |  | 0.008 | 0.014 | -0.001 | 0.013 | -0.001 | 0 | 0 | -2.9564e-05 | -63.9019 | 0 | 2.9564e-05 | 63.9019 | 0.015264 |  | 0.0002 |  | -0.003 | 0.014 | -0.002
4 | 8.19e-05 | 8.19e-05 | 0.0006 | 3.107 | -0.002 | 0.069 | 0.188 | 0.05 | -13 | 13 | -8 | 20 | 19 | 7 | 3 | 0 | 0 | 0 | 0.0007 | 0 | 1.9937 | 0.015024 |  |  | 0.01 | 0.016 | 0.002 | 0.012 | -0.002 | -0.001 | 0 | 5.9129e-05 | -27.8039 | -2.9564e-05 | 6.6108e-05 | -19.2823 | 0.01562 |  | 0.000354
5 | -8.19e-05 | -8.19e-05 | 0.0005 | 3.108 | -0.001 | 0.07 | 0.188 | 0.05 | -12 | 11 | -7 | -18 | 17 | 6 | 2 | -1 | 0 | 0 | 0.0006 | 0 | 0.99684 | 0.3155 |  |  | 0.007 | 0.013 | -0.003 | 0.012 | -0.002 | 0 | 0 | -8.8693e-05 | 8.2942 | 0 | 8.8693e-05 | -8.2942 | 0.013892 |  | 0.000251
| abs.mode | 0 |  | 3.119 | 0 | 0.074 | 0.168 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | 41.6588 |  |  | 19.5067
| Ry offsets | 8.19e-05 | 0.0001 | 3.118 | -0.001 | 0.075 | 0.169 | 0.05 | -1 | 1 | -1 | -2 | 2 | 0 | 1 | 0 | 0 | 0 | 0.0001 | 0 | 0.99684 | 0.3155
|  | 8.19e-05 | 0.0002 | 3.117 | -0.002 | 0.075 | 0.173 | 0.05 | -3 | 3 | -2 | -5 | 4 | 0 | 2 | 0 | 0 | 0 | 0.0001 | 0 | 1.9937 | 0.3155
|  | -8.19e-05 | 0.0001 | 3.118 | -0.001 | 0.075 | 0.173 | 0.05 | -2 | 1 | -1 | -2 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0.99684 | 0.3155
|  | 8.19e-05 | 0.0002 | 3.117 | -0.002 | 0.075 | 0.173 | 0.05 | -3 | 3 | -2 | -5 | 4 | 0 | 3 | 0 | 0 | 0 | 0.0001 | 0 | 1.9937 | 0.015024
|  | -8.19e-05 | 0.0001 | 3.118 | -0.001 | 0.075 | 0.173 | 0.05 | -2 | 1 | -1 | -3 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0.99684 | 0.3155 | res offset |  |  | 8.19e-05 |  |  | deg
| zero off.mode | -8.19e-05 | 0 | 3.119 | 0 | 0.075 | 0.171 | 0.05 | 0 | 0 | 0 | 0 | -1 | -1 | 1 | -1 | 0 | 0 | 0 | 0 |  |  | 5x= |  |  | 0.0004095
| abs.mode | Set to 0 | 0 | 3.119 | 0 | 0.075 | 0.171 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | 2x= |  |  | 0.0001638
[21 empty] | av.error%= | 24.7855 | 1x= |  |  | 8.19e-05
[21 empty] | stdev%= | 27.0716

## ResRz COR -2758400um
COR= -2758400um (pivot point at L1S1) | [19 empty] | expected
| Commanded | EUI | (60E-5 deg<=> 9.1564336um of Rz gauge) |  |  |  |  | EUI Hexapod Strut position |  |  | e.u (um) |  |  |  |  | hexapod EUI display |  |  |  | gauge | error
| Offset Rz | Rz value | Rz Raw | Rz Value | Y | Zx+ | Zy- | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | um | %
5x | (5x 60E-5) | 0 | 0.171 |  | 0.076 | 3.119 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.003 | 0.003 | 0.218 | 0.047 | 0.075 | 3.12 | 0.049 | 21 | -21 | 20 | -21 | 20 | -21 | 0 | 0 | 0 | 0 | 0 | 0.003 | 45.7822 | 2.6602
| 0.003 | 0.0061 | 0.265 | 0.094 | 0.077 | 3.121 | 0.047 | 41 | -42 | 41 | -41 | 41 | -42 | 0 | 0 | 0 | 0 | 0 | 0.0061 | 91.5643 | 5.3204
| -0.003 | 0.003 | 0.212 | 0.041 | 0.078 | 3.117 | 0.048 | 21 | -21 | 20 | -21 | 20 | -21 | -1 | 0 | 0 | 0 | 0 | 0.003 | 45.7822 | 10.446
| 0.003 | 0.0061 | 0.264 | 0.093 | 0.078 | 3.121 | 0.047 | 42 | -42 | 41 | -42 | 41 | -42 | -1 | 0 | 0 | 0 | 0 | 0.0061 | 91.5643 | 3.136
| -0.003 | 0.003 | 0.211 | 0.04 | 0.078 | 3.117 | 0.048 | 21 | -21 | 20 | -21 | 20 | -21 | -1 | -1 | 1 | 0 | 0 | 0.003 | 45.7822 | 12.6303
2x | (2x 60E-5) |  | 0.211 |  | 0.078 | 3.117 | 0.048 | 21 | -21 | 20 | -21 | 20 | -21 | -1 | -1 | 1 | 0 | 0 | 0.003
| 0.0012 | 0.0043 | 0.236 | 0.025 | 0.078 | 3.119 | 0.047 | 29 | -30 | 28 | -29 | 28 | -29 | -1 | -1 | 1 | 0 | 0 | 0.0043 | 18.3129 | 36.5218
| 0.0012 | 0.0055 | 0.255 | 0.044 | 0.077 | 3.12 | 0.048 | 38 | -38 | 37 | -38 | 37 | -38 | -1 | 0 | 1 | 0 | 0 | 0.0055 | 36.6257 | 40.2745
| -0.0012 | 0.0043 | 0.231 | 0.02 | 0.078 | 3.118 | 0.048 | 29 | -29 | 28 | -30 | 28 | -30 | -1 | -1 | 1 | 0 | 0 | 0.0043 | 18.3129 | 9.2143
| 0.0012 | 0.055 | 0.255 | 0.044 | 0.077 | 3.12 | 0.048 | 38 | -38 | 36 | -38 | 37 | -38 | -1 | -1 | 1 | 0 | 0 | 0.0055 | 36.6257 | 40.2745
| -0.0012 | 0.0043 | 0.231 | 0.02 | 0.078 | 3.118 | 0.048 | 29 | -29 | 28 | -30 | 28 | -30 | -1 | -1 | 1 | 0 | 0 | 0.0043 | 18.3129 | 9.2143
1x | (1x 60E-5) |  | 0.231 |  | 0.078 | 3.118 | 0.048 | 29 | -29 | 28 | -30 | 28 | -30 | -1 | -1 | 1 | 0 | 0 | 0.0043
| 0.0006 | 0.0049 | 0.245 | 0.014 | 0.078 | 3.119 | 0.048 | 33 | -34 | 32 | -34 | 32 | -34 | -1 | -1 | 1 | 0 | 0 | 0.0049 | 9.1564 | 52.8774
| 0.0006 | 0.0055 | 0.255 | 0.024 | 0.077 | 3.12 | 0.048 | 38 | -38 | 36 | -38 | 37 | -39 | -1 | 0 | 1 | 0 | 0 | 0.0055 | 18.3129 | 62.0866
| -0.0006 | 0.0049 | 0.243 | 0.012 | 0.078 | 3.118 | 0.048 | 33 | -34 | 32 | -34 | 32 | -34 | -1 | -1 | 1 | 0 | 0 | 0.0049 | 9.1564 | 31.0433
| 0.0006 | 0.0055 | 0.254 | 0.023 | 0.078 | 3.12 | 0.048 | 38 | -38 | 36 | -38 | 37 | -39 | -1 | 0 | 1 | 0 | 0 | 0.0055 | 18.3129 | 51.1696
| -0.0006 | 0.0005 | 0.243 | 0.012 | 0.078 | 3.118 | 0.048 | 33 | -34 | 32 | -34 | 32 | -34 | -1 | -1 | 1 | 0 | 0 | 0.0049 | 9.1564 | 31.0433
in abs.mode | 0 |  | 0.165 |  | 0.078 | 3.115 | 0.049 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
| 0.0006 | 0.0006 | 0.179 | 0.014 | 0.078 | 3.118 | 0.049 | 4 | -4 | 4 | -4 | 4 | -4 | 0 | 0 | 0 | 0 | 0 | 0.0006 | 9.1564 | 52.8774
| 0.0006 | 0.0012 | 0.19 | 0.025 | 0.078 | 3.118 | 0.048 | 8 | -9 | 8 | -8 | 8 | -9 | 0 | 0 | 0 | 0 | 0 | 0.0012 | 18.3129 | 73.0036
| -0.0006 | 0.0006 | 0.177 | 0.012 | 0.078 | 3.116 | 0.048 | 4 | -4 | 4 | -4 | 4 | -4 | 0 | 0 | 0 | 0 | 0 | 0.0006 | 9.1564 | 31.0433
| 0.0006 | 0.0012 | 0.189 | 0.024 | 0.078 | 3.118 | 0.048 | 8 | -9 | 8 | -8 | 8 | -8 | 0 | 0 | 0 | 0 | 0 | 0.0012 | 18.3129 | 62.0866
| -0.0006 | 0.0006 | 0.177 | 0.012 | 0.078 | 3.116 | 0.048 | 4 | -4 | 4 | -4 | 4 | -4 | 0 | 0 | 0 | 0 | 0 | 0.0006 | 9.1564 | 31.0433
0 by offset | -0.0006 | 0 | 0.166 | 0.001 | 0.077 | 3.114 | 0.049 | -1 | 0 | 0 | 0 | -1 | 1 | 0 | -1 | 0 | 0 | 0 | -0.0001
in abs.mode | Set to 0 | 0 | 0.166 | 0.001 | 0.077 | 3.114 | 0.049 | 0 | 0 | 0 | 0 | 0 | 1 | -1 | 0 | 0 | 0 | 0 | 0
[20 empty] | av.error%= | 32.3983
[20 empty] | stdev%= | 21.9721

## Repeat.RXRYRZ -2758400um
| spec= 4times the resolution for each axis |  |  |  | repetition step= |  | 0.0015° |  |  | measured w/Mitutoyo 2021 | [9 empty] | expected | hexapod |  |  | X and Y displacement due to Ry and Rx rotation respectively | [7 empty] | measured w/Mitutoyo Feb.2022
| COR= -2758400um (Pivot point at L1S1) | [18 empty] | linear | error um |  |  | .0015deg | at L1S1=> | Δ= | 0.050737 | mm= | 50.74um |  |  | measured w/Mitutoyo Feb.2022
| Commanded | EUI |  | Linear Gauges names |  |  |  | EUI Hexapod Strut position e.u. (um) | [6 empty] | hexapod EUI display (um) and (Deg) |  |  |  |  | gauge | linear gauge | [9 empty] | raw data |  |  | offset removed |  |  | measured rotations
rep. step | moves Deg | values | (+0.0015 deg<=> -50.74um of Y Gauge) | [8 empty] | spec | error< | 32.8E-5°==> |  | Y = | 0.011081 | mm= | 11.08um | value | ¦meas.-set¦ | [9 empty] | mm |  |  | mm | [5 empty] | error | x-talk
0.0015° | Rx | Rx | Zy- | Yraw | Y | Zx+ | X(Rz) | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | Y um | Rx | [8 empty] | X | y | z | X | y | z | Rx | Rx | Ry | Rz | Deg | Deg
| 0 | 0 | 0.047 | 0.08 |  | 3.115 | 0.169 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | spec < 11um |  | error< | 32.8E-5° | ===> | Error ΔY= | 0.011081 | mm= | 11.08um | -0.003 | 0.016 | 0.002 | 0 | 0 | 0 | 0 | 0
| 0.0015 | 0.0015 | 0.028 | 0.033 | -0.047 | 3.112 | 0.165 | -31 | -31 | 36 | -5 | -5 | 37 | 0 | 1 | 0 | 0.0015 | 0 | 0 | -50.7368 | 3.7368 | [8 empty] | -0.007 | -0.036 | -0.027 | -0.004 | -0.052 | -0.029 | 0.0015373 | 0.002014 |  |  | 3.7344e-05
| 0 | 0 | 0.048 | 0.08 | 0 | 3.116 | 0.169 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | [8 empty] | -0.003 | 0.016 | 0.002 | 0 | 0 | 0 | 0 | 0 |  |  | 0
| -0.0015 | -0.0015 | 0.067 | 0.13 | 0.05 | 3.117 | 0.17 | 31 | 31 | -37 | 5 | 5 | -37 | 0 | -1 | 0 | -0.0015 | 0 | 0 | 50.7368 | 0.73684 | [8 empty] | -0.006 | 0.065 | 0.021 | -0.003 | 0.049 | 0.019 | -0.0014487 | -0.0013195 |  |  | 5.1348e-05
| 0 | 0 | 0.046 | 0.082 | 0.002 | 3.111 | 0.168 | 0 | 0 | 0 | 0 | 0 | 1 | -1 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | [8 empty] | -0.008 | 0.017 | -0.002 | -0.005 | 0.001 | -0.004 | -2.9564e-05 | 0.0002778 |  |  | -2.9564e-05
| 0.0015 | 0.0015 | 0.027 | 0.034 | -0.046 | 3.111 | 0.166 | -31 | -31 | 36 | -5 | -5 | 37 | -1 | 1 | 0 | 0.0015 | 0 | 0 | -50.7368 | 4.7368 | [8 empty] | -0.009 | -0.035 | -0.028 | -0.006 | -0.051 | -0.03 | 0.0015078 | 0.0020835 |  |  | 7.7801e-06
| 0 | 0 | 0.048 | 0.082 | 0.002 | 3.115 | 0.17 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 2 | [8 empty] | -0.003 | 0.018 | 0.001 | 0 | 0.002 | -0.001 | -5.9129e-05 | 6.9449e-05 |  |  | -5.9129e-05
| -0.0015 | -0.0015 | 0.067 | 0.13 | 0.05 | 3.117 | 0.17 | 31 | 31 | -37 | 5 | 5 | -37 | 1 | -1 | 0 | -0.0015 | 0 | 0 | 50.7368 | 0.73684 | [8 empty] | -0.006 | 0.065 | 0.022 | -0.003 | 0.049 | 0.02 | -0.0014487 | -0.001389 |  |  | 5.1348e-05
| 0 | 0 | 0.046 | 0.083 | 0.003 | 3.11 | 0.167 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.008 | 0.017 | -0.002 | -0.005 | 0.001 | -0.004 | -2.9564e-05 | 0.0002778 |  |  | -2.9564e-05
| 0.0015 | 0.0015 | 0.027 | 0.035 | -0.045 | 3.111 | 0.166 | -31 | -31 | 36 | -5 | -5 | 37 | -1 | 1 | 0 | 0.0015 | 0 | 0 | -50.7368 | 5.7368 | [8 empty] | -0.009 | -0.035 | -0.027 | -0.006 | -0.051 | -0.029 | 0.0015078 | 0.002014 |  |  | 7.7801e-06
| 0 | 0 | 0.048 | 0.083 | 0.003 | 3.115 | 0.168 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.003 | 0.018 | 0.002 | 0 | 0.002 | 0 | -5.9129e-05 | 0 |  |  | -5.9129e-05
| -0.0015 | -0.0015 | 0.067 | 0.131 | 0.051 | 3.117 | 0.17 | 31 | 31 | -37 | 5 | 5 | -37 | 0 | -1 | 0 | -0.0015 | 0 | 0 | 50.7368 | 0.26316 | [8 empty] | -0.006 | 0.064 | 0.023 | -0.003 | 0.048 | 0.021 | -0.0014191 | -0.0014584 |  |  | 8.0913e-05
| 0 | 0 | 0.046 | 0.083 | 0.003 | 3.11 | 0.168 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.009 | 0.016 | -0.001 | -0.006 | 0 | -0.003 | 0 | 0.00020835 |  |  | 0
|  |  | (+0.0015 deg<=> +50.74um of X Gauge |  |  |  |  | spec 5: spec; error< 6: error< | 32.8E-5°==> |  | X = | 0.011081 | mm= | 11.08um | [16 empty] | 0.0014782 | 0.0017131 |  |  | -2.1784e-05
| Ry | Ry | X(Rz) | Y | Zy- | Xraw | X | [12 empty] | X um | Ry
| 0 | 0 | 3.099 | 0.064 | 0.028 | 0.165 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | spec < 11um |  | error< | 32.8E-5° | ===> | Error ΔX= | 0.011081 | mm= | 11.08um | -0.009 | 0.016 | -0.001 | 0 | 0 | 0 |  |  | 0
| 0.0015 | 0.0015 | 3.076 | 0.058 | 0.028 | 0.21 | 0.045 | -25 | 25 | -15 | -39 | 39 | 15 | 0 | 0 | 0 | 0 | 0.0015 | 0 | 50.7368 | 5.7368 | [8 empty] | 0.027 | 0.014 | -0.003 | 0.036 | -0.002 | -0.002 |  |  | 0.0013503 |  | -0.00014966
| 0 | 0 | 3.101 | 0.063 | 0.029 | 0.166 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | [8 empty] | -0.001 | 0.02 | 0.001 | 0.008 | 0.004 | 0.002 |  |  | 0.00030008 |  | 0.00030008
| -0.0015 | -0.0015 | 3.123 | 0.063 | 0.03 | 0.118 | -0.047 | 24 | -25 | 15 | 39 | -39 | -15 | 0 | 0 | 0 | 0 | -0.0015 | 0 | -50.7368 | 3.7368 | [8 empty] | -0.038 | 0.019 | 0.001 | -0.029 | 0.003 | 0.002 |  |  | -0.0010878 |  | 0.00041223
| 0 | 0 | 3.098 | 0.056 | 0.03 | 0.163 | -0.002 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | [8 empty] | -0.01 | 0.017 | 0 | -0.001 | 0.001 | 0.001 |  |  | -3.7509e-05 |  | -3.7509e-05
| 0.0015 | 0.0015 | 3.077 | 0.057 | 0.03 | 0.21 | 0.045 | -25 | 25 | -15 | -39 | 39 | 15 | 0 | 0 | 0 | 0 | 0.0015 | 0 | 50.7368 | 5.7368 | [8 empty] | 0.027 | 0.015 | -0.003 | 0.036 | -0.001 | -0.002 |  |  | 0.0013503 |  | -0.00014966
| 0 | 0 | 3.102 | 0.062 | 0.029 | 0.165 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [8 empty] | -0.002 | 0.02 | 0 | 0.007 | 0.004 | 0.001 |  |  | 0.00026257 |  | 0.00026257
| -0.0015 | -0.0015 | 3.123 | 0.062 | 0.031 | 0.118 | -0.047 | 24 | -25 | 15 | 39 | -39 | -15 | 0 | 0 | 0 | 0 | -0.0015 | 0 | -50.7368 | 3.7368 | [8 empty] | -0.041 | 0.02 | 0.001 | -0.032 | 0.004 | 0.002 |  |  | -0.0012003 |  | 0.0002997
| 0 | 0 | 3.099 | 0.056 | 0.031 | 0.162 | -0.003 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.01 | 0.017 | -0.001 | -0.001 | 0.001 | 0 |  |  | -3.7509e-05 |  | -3.7509e-05
| 0.0015 | 0.0015 | 3.077 | 0.057 | 0.03 | 0.21 | 0.045 | -25 | 25 | -15 | -39 | 39 | 15 | 0 | -1 | 0 | 0 | 0.0015 | 0 | 50.7368 | 5.7368 | [8 empty] | 0.026 | 0.014 | -0.003 | 0.035 | -0.002 | -0.002 |  |  | 0.0013128 |  | -0.00018717
| 0 | 0 | 3.102 | 0.062 | 0.03 | 0.165 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [8 empty] | -0.003 | 0.02 | 0.001 | 0.006 | 0.004 | 0.002 |  |  | 0.00022506 |  | 0.00022506
| -0.0015 | -0.0015 | 3.124 | 0.062 | 0.032 | 0.117 | -0.048 | 24 | -25 | 15 | 39 | -39 | -15 | 0 | 0 | 0 | 0 | -0.0015 | 0 | -50.7368 | 2.7368 | [8 empty] | -0.038 | 0.02 | 0 | -0.029 | 0.004 | 0.001 |  |  | -0.0010878 |  | 0.00041223
| 0 | 0 | 3.099 | 0.057 | 0.031 | 0.162 | -0.003 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.01 | 0.017 | 0.001 | -0.001 | 0.001 | 0.002 |  |  | -3.7509e-05 |  | -3.7509e-05
|  |  | (+0.0015 deg<=> -22.89um of X(Rz) Gauge) |  |  |  |  | spec 5: spec; error< 6: error< | 240E-5°==> |  | X(Rz)= | 0.036626 | mm= | 36.63um | [18 empty] | 0.0012316 |  | -0.00026844
| Rz | Rz | X(Rz)raw | X(Rz) | Y | Zy- | Zx+ | [12 empty] | X(Rz) um | Rz
| 0 | 0 | 0.163 |  | 0.057 | 0.031 | 3.099 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | spec <37um |  | error< | 240E-5° | ===> | error Rz gauge= |  | 0.036626 | 36.63um | -0.01 | 0.017 | 0.001 | 0 | 0 | 0 |  |  |  | 0
| 0.0015 | 0.0015 | 0.189 | 0.026 | 0.058 | 0.031 | 3.101 | 11 | -11 | 11 | -10 | 10 | -11 | 0 | 0 | 0 | 0 | 0 | 0.0015 | 22.8911 | 3.1089 | [8 empty] | 0.015 | 0.018 | 0 | 0.025 | 0.001 | -0.001 |  |  |  | 0.0018482 | 0.00034825
| 0 | 0 | 0.16 | -0.003 | 0.06 | 0.031 | 3.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | [8 empty] | -0.01 | 0.018 | 0 | 0 | 0.001 | -0.001 |  |  |  | 0 | 0
| -0.0015 | -0.0015 | 0.137 | -0.026 | 0.06 | 0.032 | 3.099 | -11 | 11 | -10 | 10 | -11 | 11 | 0 | -1 | 0 | 0 | 0 | -0.0015 | -22.8911 | 3.1089 | [8 empty] | -0.029 | 0.016 | 0 | -0.019 | -0.001 | -0.001 |  |  |  | -0.0014047 | 9.5333e-05
| 0 | 0 | 0.167 | 0.004 | 0.061 | 0.031 | 3.102 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | [8 empty] | -0.004 | 0.018 | 0.001 | 0.006 | 0.001 | 0 |  |  |  | 0.00044358 | 0.00044358
| 0.0015 | 0.0015 | 0.19 | 0.027 | 0.059 | 0.032 | 3.102 | 11 | -10 | 10 | -10 | 10 | -11 | 0 | 0 | 0 | 0 | 0 | 0.0015 | 22.8911 | 4.1089 | 9:37 AM, Faulted: Following Errror, Drive Fault, Simulink Fault. Clear error Command | [7 empty] | 0.015 | 0.019 | 0.001 | 0.025 | 0.002 | 0 |  |  |  | 0.0018482 | 0.00034825
| 0 | 0 | 0.161 | -0.002 | 0.061 | 0.032 | 3.101 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 9:39 AM, Faulted: Drive Fauls, Simulink Fault. Clear Error Command | [7 empty] | -0.011 | 0.018 | 0.001 | -0.001 | 0.001 | 0 |  |  |  | -7.393e-05 | -7.393e-05
| -0.0015 | -0.0015 | 0.137 | -0.026 | 0.062 | 0.033 | 3.1 | -11 | 11 | -11 | 10 | -11 | 11 | 0 | 0 | 0 | 0 | 0 | -0.0015 | -22.8911 | 3.1089 | [8 empty] | -0.03 | 0.016 | 0 | -0.02 | -0.001 | -0.001 |  |  |  | -0.0014786 | 2.1403e-05
| 0 | 0 | 0.167 | 0.004 | 0.062 | 0.032 | 3.102 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | [8 empty] | -0.004 | 0.018 | 0.001 | 0.006 | 0.001 | 0 |  |  |  | 0.00044358 | 0.00044358
| 0.0015 | 0.0015 | 0.19 | 0.027 | 0.06 | 0.032 | 3.102 | 11 | -10 | 10 | -10 | 10 | -11 | 0 | 0 | 0 | 0 | 0 | 0.0015 | 22.8911 | 4.1089 | [8 empty] | 0.015 | 0.019 | 0 | 0.025 | 0.002 | -0.001 |  |  |  | 0.0018482 | 0.00034825
| 0 | 0 | 0.161 | -0.002 | 0.061 | 0.033 | 3.102 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | [8 empty] | -0.01 | 0.019 | 0.001 | 0 | 0.002 | 0 |  |  |  | 0 | 0
| -0.0015 | -0.0015 | 0.136 | -0.027 | 0.062 | 0.034 | 3.1 | -11 | 11 | -11 | 10 | -11 | 11 | 0 | -1 | 0 | 0 | 0 | -0.0015 | -22.8911 | 4.1089 |  |  | X(Rz) gauge displacement measurement due to Rz rotation | [5 empty] | -0.028 | 0.017 | 0 | -0.018 | 0 | -0.001 |  |  |  | -0.0013307 | 0.00016926
| 0 | 0 | 0.166 | 0.003 | 0.062 | 0.033 | 3.103 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 |  |  | .0015deg | at L1S1=> | Δ= | 0.022891 | mm= | 22.89um | -0.004 | 0.019 | 0.001 | 0.006 | 0.002 | 0 |  |  |  | 0.00044358 | 0.00044358
[39 empty] | 0.0016265 | 0.00012646
0
X
Y
Z

## RepeatXYZ COR -2758400um
3.3.12 Hexapod Repeatability | [21 empty] | measured w/mitutoyo Feb.2022
| spec= 4 times resolution of each axis | [18 empty] | hexapod
| COR= -2758400um (Pivot point at L1S1) |  |  | Moving 50um in absolute mode | [15 empty] | error |  | measured w/mitutoyo Feb.2022
| Commanded | EUI | [5 empty] | EUI Hexapod Strut position e.u. (um) | [6 empty] | hexapod EUI display (um) and (Deg) |  |  |  |  | ¦meas.-set¦ |  | raw data |  |  | offset removed
| moves um | values | Gauge X(Rz) |  |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | X | Y | Z | Rx | Ry | Rz | um |  | mm |  |  | mm
| X Move | X um | X Raw | X Value | Gauge Y | Gauge Zx+ | Gauge Zy- | [12 empty] | X | x | y | z | x | y | z | X error | x-talk
| 0 | 0 | 0.291 | 0 | 0.044 | 3.115 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | spec < 20um | -0.004 | 0.015 | 0.002 | 0 | 0 | 0 | (um) | (um)
| 50 | 50 | 0.341 | 0.05 | 0.043 | 3.116 | 0.049 | -25 | 25 | 0 | -25 | 25 | 0 | 51 | 0 | 0 | 0 | 0 | 0 | 4.2633e-14 | 0.046 | 0.015 | 0.001 | 0.05 | 0 | -0.001 | 0 | 0.70711
| 0 | -1 | 0.291 | 0 | 0.046 | 3.116 | 0.049 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.005 | 0.014 | 0.001 | -0.001 | -0.001 | -0.001 | -1 | 6.1332e-16
| -50 | -51 | 0.243 | -0.048 | 0.045 | 3.115 | 0.051 | 25 | -25 | 0 | 25 | -25 | 0 | -51 | 0 | 0 | 0 | 0 | 0 | 3 | -0.055 | 0.013 | 0 | -0.051 | -0.002 | -0.002 | -1 | 0
| 0 | 0 | 0.292 | 0.001 | 0.041 | 3.115 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | -0.003 | 0.015 | 0.002 | 0.001 | 0 | 0 | 1 | 0
| 50 | 50 | 0.341 | 0.05 | 0.043 | 3.116 | 0.049 | -25 | 25 | 0 | -25 | 25 | 0 | 51 | 0 | 0 | 0 | 0 | 0 | 4.2633e-14 | 0.046 | 0.015 | 0 | 0.05 | 0 | -0.002 | 0 | 1.4142
| 0 | -1 | 0.291 | 0 | 0.046 | 3.116 | 0.049 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.005 | 0.015 | 0.001 | -0.001 | 0 | -0.001 | -1 | 0.70711
| -50 | -51 | 0.243 | -0.048 | 0.045 | 3.115 | 0.051 | 25 | -25 | 0 | 25 | -25 | 0 | -51 | 0 | 0 | 0 | 0 | 0 | 3 | -0.055 | 0.014 | 0 | -0.051 | -0.001 | -0.002 | -1 | 0.70711
| 0 | 0 | 0.292 | 0.001 | 0.041 | 3.115 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | -0.003 | 0.015 | 0.003 | 0.001 | 0 | 0.001 | 1 | 0.70711
| 50 | 50 | 0.341 | 0.05 | 0.042 | 3.116 | 0.048 | -25 | 25 | 0 | -25 | 25 | 0 | 51 | 0 | 0 | 0 | 0 | 0 | 4.2633e-14 | 0.047 | 0.015 | 0.001 | 0.051 | 0 | -0.001 | 1 | 0.70711
| 0 | -1 | 0.291 | 0 | 0.045 | 3.116 | 0.049 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.005 | 0.014 | 0.002 | -0.001 | -0.001 | 0 | -1 | 0.70711
| -50 | -51 | 0.243 | -0.048 | 0.045 | 3.115 | 0.051 | 25 | -25 | 0 | 25 | -25 | 0 | -51 | 0 | 0 | 0 | 0 | 0 | 3 | -0.055 | 0.013 | 0.001 | -0.051 | -0.002 | -0.001 | -1 | 0.70711
| 0 | 0 | 0.292 | 0.001 | 0.041 | 3.115 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | -0.004 | 0.015 | 0.003 | 0 | 0 | 0.001 | 0 | 0.70711 | 0.86603 | 0.58926
| Y Move | Y um | Y Raw | Y Value | Gauge X | Gauge Zx+ | Gauge Zy- | [12 empty] | Y | [6 empty] | y error | x-talk
| 0 | 0 | 0.041 | 0 | 0.292 | 3.115 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | spec < 20um | -0.004 | 0.015 | 0.003 | 0 | 0 | 0 | (um) | (um)
| 50 | 51 | 0.091 | 0.05 | 0.3 | 3.121 | 0.051 | 14 | 15 | -29 | 15 | 15 | -29 | 1 | 51 | 0 | 0 | 0 | 0 | 1 | -0.002 | 0.064 | 0.002 | 0.002 | 0.049 | -0.001 | -1 | 2.1213
| 0 | -1 | 0.041 | 0 | 0.291 | 3.114 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | -1 | 0 | 0 | 0 | 0 | 1 | -0.008 | 0.013 | 0.001 | -0.004 | -0.002 | -0.002 | -2 | 1.4142
| -50 | -51 | -0.006 | -0.047 | 0.286 | 3.115 | 0.049 | -15 | -15 | 29 | -15 | -15 | 29 | -1 | -51 | 0 | 0 | 0 | 0 | 4 | -0.01 | -0.039 | -0.003 | -0.006 | -0.054 | -0.006 | -4 | 0
| 0 | 1 | 0.043 | 0.002 | 0.293 | 3.118 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | -0.002 | 0.016 | 0.002 | 0.002 | 0.001 | -0.001 | 1 | 2.1213
| 50 | 51 | 0.092 | 0.051 | 0.3 | 3.121 | 0.051 | 14 | 15 | -29 | 15 | 15 | -29 | 1 | 51 | 0 | 0 | 0 | 0 | 0 | -0.002 | 0.065 | 0.002 | 0.002 | 0.05 | -0.001 | 0 | 2.1213
| 0 | -1 | 0.042 | 0.001 | 0.291 | 3.114 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | -1 | 0 | 0 | 0 | 0 | 2 | -0.007 | 0.014 | 0.001 | -0.003 | -0.001 | -0.002 | -1 | 0.70711
| -50 | -51 | -0.006 | -0.047 | 0.286 | 3.115 | 0.049 | -15 | -15 | 29 | -15 | -15 | 29 | -1 | -51 | 0 | 0 | 0 | 0 | 4 | -0.009 | -0.038 | -0.003 | -0.005 | -0.053 | -0.006 | -3 | 0.70711
| 0 | 1 | 0.043 | 0.002 | 0.293 | 3.118 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | -0.002 | 0.016 | 0.002 | 0.002 | 0.001 | -0.001 | 1 | 2.1213
| 50 | 51 | 0.097 | 0.056 | 0.304 | 3.12 | 0.051 | 15 | 14 | -29 | 14 | 14 | -29 | 0 | 50 | 0 | 0 | 0 | 0 | 5 | -0.003 | 0.065 | 0.002 | 0.001 | 0.05 | -0.001 | 0 | 1.4142
| 0 | -1 | 0.044 | 0.003 | 0.291 | 3.114 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 4 | -0.007 | 0.014 | 0.001 | -0.003 | -0.001 | -0.002 | -1 | 0.70711
| -50 | -51 | -0.004 | -0.045 | 0.286 | 3.115 | 0.049 | -15 | -15 | 29 | -15 | -15 | 29 | -1 | -50 | 0 | 0 | 0 | 0 | 6 | -0.009 | -0.038 | -0.002 | -0.005 | -0.053 | -0.005 | -3 | 8.6736e-16
| 0 | 1 | 0.047 | 0.006 | 0.293 | 3.118 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 5 | -0.002 | 0.016 | 0.002 | 0.002 | 0.001 | -0.001 | 1 | 2.1213 | 1.7056 | 1.2964
| Z Move | Z um | Z Raw | Z Value | Gauge X | Gauge Zx+ | Gauge Y | [12 empty] | Z | [6 empty] | z error | x-talk
| 0 | 0 | 0.047 | 0 | 0.293 | 3.116 | 0.046 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | spec < 4um | -0.002 | 0.016 | 0.002 | 0 | 0 | 0 | (um) | (um)
| 50 | 50 | 0.094 | 0.047 | 0.291 | 3.169 | 0.046 | -41 | -41 | -41 | -41 | -41 | -41 | 1 | 0 | 50 | 0 | 0 | 0 | 3 | -0.01 | 0.012 | 0.051 | -0.008 | -0.004 | 0.049 | -1 | 2.8284
| 0 | 0 | 0.046 | -0.001 | 0.29 | 3.114 | 0.047 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.01 | 0.012 | 0.004 | -0.008 | -0.004 | 0.002 | 2 | 2.8284
| -50 | -50 | 0 | -0.047 | 0.29 | 3.063 | 0.046 | 41 | 41 | 41 | 41 | 41 | 41 | -1 | 0 | -50 | 0 | 0 | 0 | 3 | -0.002 | 0.01 | -0.045 | 0 | -0.006 | -0.047 | 3 | 4.2426
| 0 | 0 | 0.048 | 0.001 | 0.292 | 3.116 | 0.045 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.004 | 0.013 | 0.004 | -0.002 | -0.003 | 0.002 | 2 | 0.70711
| 50 | 50 | 0.094 | 0.047 | 0.292 | 3.168 | 0.045 | -41 | -41 | -41 | -41 | -41 | -41 | 1 | 0 | 50 | 0 | 0 | 0 | 3 | -0.01 | 0.013 | 0.051 | -0.008 | -0.003 | 0.049 | -1 | 3.5355
| 0 | 0 | 0.046 | -0.001 | 0.29 | 3.114 | 0.047 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.009 | 0.014 | 0.002 | -0.007 | -0.002 | 0 | 0 | 3.5355
| -50 | -50 | 0 | -0.047 | 0.291 | 3.063 | 0.046 | 41 | 41 | 41 | 41 | 41 | 41 | -1 | 0 | -50 | 0 | 0 | 0 | 3 | -0.002 | 0.012 | -0.046 | 0 | -0.004 | -0.048 | 2 | 2.8284
| 0 | 0 | 0.048 | 0.001 | 0.292 | 3.116 | 0.044 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | -0.003 | 0.014 | 0.003 | -0.001 | -0.002 | 0.001 | 1 | 0.70711
| 50 | 50 | 0.094 | 0.047 | 0.292 | 3.168 | 0.045 | -41 | -41 | -41 | -41 | -41 | -41 | 1 | 0 | 50 | 0 | 0 | 0 | 3 | -0.01 | 0.013 | 0.051 | -0.008 | -0.003 | 0.049 | -1 | 3.5355
| 0 | 0 | 0.046 | -0.001 | 0.291 | 3.114 | 0.047 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.01 | 0.014 | 0.002 | -0.008 | -0.002 | 0 | 0 | 4.2426
| -50 | -50 | 0 | -0.047 | 0.291 | 3.062 | 0.046 | 41 | 41 | 41 | 41 | 41 | 41 | -1 | 0 | -50 | 0 | 0 | 0 | 3 | -0.002 | 0.011 | -0.046 | 0 | -0.005 | -0.048 | 2 | 3.5355
| 0 | 0 | 0.048 | 0.001 | 0.292 | 3.116 | 0.045 | 0 | 0 | 0 | 0 | 0 | -1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | -0.003 | 0.015 | 0.003 | -0.001 | -0.001 | 0.001 | 1 | 6.1332e-16 | 1.4035 | 2.7106

## COR from March and June2021
Data June 2021 after fixing hexapod
3.3.2 | March2021 |  | limited movement because actuator 3 faulted | [12 empty] | Rx | Y
| Camera Hexapod set values (mm and Deg) |  |  |  |  | Laser tracker masured values |  |  | measured | predicted | error | error |  |  |  | 0.24 | -7.98
X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Angle°¦ | Y (=Z*Rx) | mm | mm |  |  |  | 0.17 | -5.6175
X: w/COR=-820.4mm (Z=0mm); Laser tracker masured values Y: mm | plane to plane | mm |  | (offset removed) |  | top | low | -0.17 | 5.6825
0 | 0 | 0 | 0.04 | 0 | 0 | 0.11 | 0.06 | 0.09 | 0.0375 | 0.023562 | 0.036438 | -0.030562 | 0.00093404 | 0.12 | -0.12
0 | 0 | 0 | 0.12 | 0 | 0 | 0.13 | 0.14 | 0.08 | 0.1177 | 0.070686 | 0.069314 | 0.002314 | 5.3546e-06 | 0.12 | -0.12 | -0.24 | 7.99
0 | 0 | 0 | 0.2 | 0 | 0 | 0.09 | 0.14 | 0.09 | 0.1959 | 0.11781 | 0.02219 | -0.04481 | 0.0020079 | 0.12 | -0.12
0 | 0 | 0 | 0.24 | 0 | 0 | 0.1 | 0.16 | 0.08 | 0.2368 | 0.14137 | 0.018628 | -0.048372 | 0.0023399 | 0.12 | -0.12
0 | 0 | 0 | -0.08 | 0 | 0 | 0.09 | 0.08 | 0.08 | 0.0832 | -0.047124 | 0.12712 | 0.060124 | 0.0036149 | 0.12 | -0.12
0 | 0 | 0 | -0.17 | 0 | 0 | 0.09 | 0.03 | 0.08 | 0.1738 | -0.10014 | 0.13014 | 0.063138 | 0.0039865 | 0.12 | -0.12
Camera Hexapod set values (mm and Deg) Rx: -0.24 |  | -0.14137 |  |  |  | 0.12 | -0.12
X: w/COR =-2758.4mm (Z=-1938mm) |  | offset = | 0.067305
0 | 0 | 0 | 0.08 | 0 | 0 | 0.08 | -2.56 | 0.07 | 0.0764 | -2.6588 | 0.098841 | 0.045841 | 0.0021014 | 0.12 | -0.12
0 | 0 | 0 | 0.16 | 0 | 0 | 0.06 | -5.24 | 0.08 | 0.1562 | -5.3177 | 0.077682 | 0.024682 | 0.00060918 | 0.12 | -0.12
Camera Hexapod set values (mm and Deg) Rx: 0.17; Laser tracker masured values Y: -5.6175 |  | -5.65 | 0.032537 | -0.020463 | 0.00041875 | 0.12 | -0.12
0 | 0 | 0 | 0.2 | 0 | 0 | 0.06 | -6.58 | 0.06 | 0.1939 | -6.6471 | 0.067102 | 0.014102 | 0.00019887 | 0.12 | -0.12
Camera Hexapod set values (mm and Deg) Rx: 0.24; Laser tracker masured values Y: -7.98 |  | -7.9765 | -0.0034776 | -0.056478 | 0.0031897 | 0.12 | -0.12
0 | 0 | 0 | -0.04 | 0 | 0 | 0.1 | 1.41 | 0.08 | 0.0444 | 1.3294 | 0.08058 | 0.02758 | 0.00076063 | 0.12 | -0.12
0 | 0 | 0 | -0.08 | 0 | 0 | 0.1 | 2.74 | 0.08 | 0.0826 | 2.6588 | 0.081159 | 0.028159 | 0.00079294 | 0.12 | -0.12
Camera Hexapod set values (mm and Deg) Rx: -0.17; Laser tracker masured values Y: 5.6825 |  | 5.65 | 0.032463 | -0.020537 | 0.00042176 | 0.12 | -0.12
Camera Hexapod set values (mm and Deg) Rx: -0.24; Laser tracker masured values Y: 7.99 |  | 7.9765 | 0.013478 | -0.039522 | 0.001562 | 0.12 | -0.12
X: in red: added data from June 2021 (after fixing the hexapod actuators) |  | offset = | 0.053374 | 0.041465
[11 empty] | error RMS= | 0.045671
Camera Hexapod set values (mm and Deg) Y: Camera Hexapod set values (mm and Deg); Laser tracker masured values X: Laser tracker masured values | measured | predicted | error
X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Angle°¦ | Y (=Z*Rx) | mm
X: w/COR=-820.4mm (Z=0mm); Laser tracker masured values Y: mm | plane to plane | mm | (offset removed)
0 | 0 | 0 | 0.04 | 0 | 0 | 0.11 | 0.06 | 0.09 | 0.0375 | 0.023562 | -0.030562
0 | 0 | 0 | 0.12 | 0 | 0 | 0.13 | 0.14 | 0.08 | 0.1177 | 0.070686 | 0.002314
0 | 0 | 0 | 0.2 | 0 | 0 | 0.09 | 0.14 | 0.09 | 0.1959 | 0.11781 | -0.04481
0 | 0 | 0 | 0.24 | 0 | 0 | 0.1 | 0.16 | 0.08 | 0.2368 | 0.14137 | -0.048372
0 | 0 | 0 | -0.08 | 0 | 0 | 0.09 | 0.08 | 0.08 | 0.0832 | -0.047124 | 0.060124
0 | 0 | 0 | -0.17 | 0 | 0 | 0.09 | 0.03 | 0.08 | 0.1738 | -0.10014 | 0.063138
w/COR =-2758.4mm (Z=-1938mm)
0 | 0 | 0 | 0.08 | 0 | 0 | 0.08 | -2.56 | 0.07 | 0.0764 | -2.6588 | 0.017841
0 | 0 | 0 | 0.16 | 0 | 0 | 0.06 | -5.24 | 0.08 | 0.1562 | -5.3177 | -0.0033184
0 | 0 | 0 | 0.2 | 0 | 0 | 0.06 | -6.58 | 0.06 | 0.1939 | -6.6471 | -0.013898
0 | 0 | 0 | -0.04 | 0 | 0 | 0.1 | 1.41 | 0.08 | 0.0444 | 1.3294 | -0.0004204
0 | 0 | 0 | -0.08 | 0 | 0 | 0.1 | 2.74 | 0.08 | 0.0826 | 2.6588 | 0.0001592
RMS= sqrt(sum[(x-x0)**2]*1/n)

## COR march 2022_comcam mounted
[5 empty] | from 2021 | Data
|  | Camera Hexapod set values (mm and Deg) |  |  |  |  | Laser tracker masured values |  |  | measured | predicted | error
| X | Y | Z | Rx | Ry | Rz | X | Y | Z | ¦Angle°¦ | Y (=Z*Rx) | mm
| X: w/COR=-820.4mm (Z=0mm); Laser tracker masured values Y: mm | plane to plane | mm | (offset removed)
| 0 | 0 | 0 | 0.04 | 0 | 0 | 0.11 | 0.06 | 0.09 | 0.0375 | 0.023562 | -0.030562
| 0 | 0 | 0 | 0.12 | 0 | 0 | 0.13 | 0.14 | 0.08 | 0.1177 | 0.070686 | 0.002314
| 0 | 0 | 0 | 0.2 | 0 | 0 | 0.09 | 0.14 | 0.09 | 0.1959 | 0.11781 | -0.04481
| 0 | 0 | 0 | 0.24 | 0 | 0 | 0.1 | 0.16 | 0.08 | 0.2368 | 0.14137 | -0.048372
| 0 | 0 | 0 | -0.08 | 0 | 0 | 0.09 | 0.08 | 0.08 | 0.0832 | -0.047124 | 0.060124
| 0 | 0 | 0 | -0.17 | 0 | 0 | 0.09 | 0.03 | 0.08 | 0.1738 | -0.10014 | 0.063138
from data 2022
LTS-206 3.3.13. Rotation Absolute Accuracy tests (assembled from positioning and maximum range data)
Rx rotational absolute accuracy test | Laser tracker masured values Z: Zero Values | [6 empty] | measured
23/3/2022 | Laser tracker masured values Y: p1; Laser tracker masured values Z: 859.0775 | -0.042503 | -19.0742 |  |  |  |  | minus
| X: COR= -2758400um; Laser tracker masured values Y: p2; Laser tracker masured values Z: -859.0758 | 0.88609 | -19.0754 |  |  |  |  | predicted
| Camera Hexapod set values (mm and Deg) Y: trans. in mm; Camera Hexapod set values (mm and Deg) Ry: rot. in Deg; Laser tracker masured values X: expected; Laser tracker masured values Y: reduced laser tracker data mm |  | measured rotation angle in Deg |  |  |  | error | error
Move# | X | Y | Z | Rx | Ry | Rz | shift | X | Y | Z |  | Rx | Ry | Rz | Deg | mm
| Camera Hexapod set values (mm and Deg) Rx: 0; Laser tracker masured values X: 0; Laser tracker masured values Y: 0.0001395; Laser tracker masured values Z: -5e-07 | -0.001435 |  | 1.4782e-08 |  |  | 1.4782e-08
25% | Camera Hexapod set values (mm and Deg) Rx: 0.06; Laser tracker masured values X: -2.0295; Laser tracker masured values Y: 0.0075875; Laser tracker masured values Z: -2.0164 | -0.00146 |  | 0.059614 |  |  | -0.00038622 | 0.013064
50 | Camera Hexapod set values (mm and Deg) Rx: 0.12; Laser tracker masured values X: -4.0589; Laser tracker masured values Y: 0.031523; Laser tracker masured values Z: -4.0415 | -0.011987 |  | 0.11948 |  |  | -0.00051561 | 0.01744
75 | Camera Hexapod set values (mm and Deg) Rx: 0.18; Laser tracker masured values X: -6.0884; Laser tracker masured values Y: 0.041216; Laser tracker masured values Z: -6.0763 | -0.029301 |  | 0.17964 |  |  | -0.00035753 | 0.012093
100 | Camera Hexapod set values (mm and Deg) Rx: 0.24; Laser tracker masured values X: -8.1179; Laser tracker masured values Y: 0.027002; Laser tracker masured values Z: -8.1081 | -0.03417 |  | 0.23971 |  |  | -0.00028813 | 0.0097459
-25 | Camera Hexapod set values (mm and Deg) Rx: -0.06; Laser tracker masured values X: 2.0295; Laser tracker masured values Y: 0.024608; Laser tracker masured values Z: 2.0395 | -0.000622 |  | -0.060297 |  |  | -0.00029697 | 0.010045
-50 | Camera Hexapod set values (mm and Deg) Rx: -0.12; Laser tracker masured values X: 4.0589; Laser tracker masured values Y: 0.01513; Laser tracker masured values Z: 4.077 | -0.012543 |  | -0.12053 |  |  | -0.00053404 | 0.018064
-75 | Camera Hexapod set values (mm and Deg) Rx: -0.18; Laser tracker masured values X: 6.0884; Laser tracker masured values Y: 0.019829; Laser tracker masured values Z: 6.1077 | -0.005701 |  | -0.18057 |  |  | -0.00056856 | 0.019231
-100 | Camera Hexapod set values (mm and Deg) Rx: -0.24; Laser tracker masured values X: 8.1179; Laser tracker masured values Y: -3.85e-05; Laser tracker masured values Z: 8.13 | -0.008541 |  | -0.24036 |  |  | -0.0003576 | 0.012096
Ry rotational absolute accuracy test | bar: Laser tracker masured values X
| Camera Hexapod set values (mm and Deg) Ry: 0; Laser tracker masured values X: 0; Laser tracker masured values Y: 6.35e-05; Laser tracker masured values Z: 0 | -1e-06 |  |  | 1.8773e-06 |  | 1.8773e-06
25% | Camera Hexapod set values (mm and Deg) Ry: 0.06; Laser tracker masured values X: 2.0295; Laser tracker masured values Y: 2.0254; Laser tracker masured values Z: 0.017499 | -0.016198 |  |  | 0.05988 |  | -0.00011962 | -0.0040461
50 | Camera Hexapod set values (mm and Deg) Ry: 0.12; Laser tracker masured values X: 4.0589; Laser tracker masured values Y: 4.0543; Laser tracker masured values Z: -0.009728 | -0.015929 |  |  | 0.11986 |  | -0.00013625 | -0.0046087
75 | Camera Hexapod set values (mm and Deg) Ry: 0.18; Laser tracker masured values X: 6.0884; Laser tracker masured values Y: 6.0811; Laser tracker masured values Z: 0.010779 | -0.025934 |  |  | 0.17978 |  | -0.00021592 | -0.0073033
100 | Camera Hexapod set values (mm and Deg) Ry: 0.24; Laser tracker masured values X: 8.1179; Laser tracker masured values Y: 8.1293; Laser tracker masured values Z: 0.0006175 | -0.033249 |  |  | 0.24034 |  | 0.00033769 | 0.011422
-25 | Camera Hexapod set values (mm and Deg) Ry: -0.06; Laser tracker masured values X: -2.0295; Laser tracker masured values Y: -2.0367; Laser tracker masured values Z: 0.018046 | -0.010976 |  |  | -0.060215 |  | -0.00021475 | -0.0072639
-50 | Camera Hexapod set values (mm and Deg) Ry: -0.12; Laser tracker masured values X: -4.0589; Laser tracker masured values Y: -4.0582; Laser tracker masured values Z: 0.026688 | 0.011606 |  |  | -0.11998 |  | 2.3465e-05 | 0.0007937
-75 | Camera Hexapod set values (mm and Deg) Ry: -0.18; Laser tracker masured values X: -6.0884; Laser tracker masured values Y: -6.0928; Laser tracker masured values Z: 0.022048 | -0.015558 |  |  | -0.18013 |  | -0.00012997 | -0.0043962
-100 | Camera Hexapod set values (mm and Deg) Ry: -0.24; Laser tracker masured values X: -8.1179; Laser tracker masured values Y: -8.1132; Laser tracker masured values Z: 0.013048 | -0.020744 |  |  | -0.23986 |  | 0.000138 | 0.0046679